## Index
EQUIPMENT DETAILS
Sl.No | Product | Machine | Description | DETAILS
1 | AirTube: Transfer Station
|  | HR-ES / LS | Take Over Station | Status bits | Positions | Commands
2 | AirTube: Sending Station
|  | HR-HSK/B | Machine for sending a sample via pneumatic tube as well as for receiving the returning carrier | Status bits | Positions | Commands
|  | HR-BM | The HR- BMK is a fully automatic plant station with mixer and compressor. | Status bits | Positions | Commands
3 | AirTube: Receiving Station
|  | HR-HSK/L | The HR-HSK/L is specially designed for sending and receiving tube capsule for robot automations. | Status bits | Positions | Commands
|  | HR-LSP | Fully automatic dispatch and receiving station for the receiving of granular and powder samples and return of the emptied capsules to the plant. | Status bits | Positions | Commands
|  | HR-LA | Fully automatic dispatch and receiving station for the receiving of granular and powder samples and return of the emptied capsules to the plant. | Status bits | Positions | Commands
4 | Milling Machine
|  | HS-FF2000 | HS-FF 2000 is the milling machine for automatic, analysis-ready preparation of iron and steel samples for spectral analysis | Status bits | Positions | Commands
|  | HS-CF | The HS-CF is the ideal machine for the fully automatic processing of cylindrical and conical steel and iron samples | Status bits | Positions | Commands
|  | HS-F 3000 | HS-FF 2000 is the milling machine for automatic, analysis-ready preparation of iron and steel samples for spectral analysis | Status bits | Positions | Commands
5 | Crushers
|  | HP-WCSA | HP-WCSA is the perfect machine for weighing, crushing and splitting. | Status bits | Positions | Commands
|  | HP-WCSA Lift |  | Status bits | Positions | Commands
|  | HP-BTM | The HP-BTM is a combined machine for crushing, splitting and grinding of coarse ores, drill cores, minerals and similar materials | Status bits | Positions | Commands
|  | HP-CA | The HP-CA is fully automatic crushing of iron ores and similar materials | Status bits | Positions | Commands
6 | Pulverizers
|  | HP-M | Fully automatic sample preparation for X-ray fluorescent analysis with all its advantages is made by HERZOG pulverizing mills and pellet presses in modular design | Status bits | Positions | Commands
|  | HP-MA | The HP-MA is suitable for grinding a wide variety of minerals such as cement raw meal, clinker, cement, slag, ores, oxydes and ferro-alloys. | Status bits | Positions | Commands
|  | HP-M1500 | The automatic machining cycle ensures reproducible and fast sample preparation | Status bits | Positions | Commands
|  | HP-SCD | The HP-SCD is dosing system suitable for cement samples | Status bits | Positions | Commands
7 | Pelletizing presses
|  | HP-PA | The HP-PA produces high quality pressed pellets with a smooth surface necessary for an optimal XRF results | Status bits | Positions | Commands
8 | Separating and stamping Machines
|  | HUST | The HUST is a sample punching machine for determination of carbon-, nitrogen- and sulfur concentration in steel. | Status bits | Positions | Commands
9 | Analyzer/Spectrometer/Scanner
|  | ARL_SMS XRF |  | Status bits | Positions | Commands
|  | SuperQ XRF |  | Status bits | Positions | Commands
|  | Sympatec Analyzer Database |  | Status bits | Positions | Commands
|  | LECO CNS Analyzer |  | Status bits | Positions | Commands
|  | LECO Manual Gas Analyser |  | Status bits | Positions | Commands
|  | ARL_SMS OES |  | Status bits | Positions | Commands
|  | OBLF OES |  | Status bits | Positions | Commands
|  | Impromat RMS |  | Status bits | Positions | Commands
|  | Symbol Barcode Scanner |  | Status bits | Positions | Commands
|  | Keyence Barcode Scanner |  | Status bits | Positions | Commands
|  | AccuPyc Density Meter |  | Status bits | Positions | Commands
10 | Magazine
|  | Input Cup Magazine (Virtual) |  | Status bits | Positions | Commands
|  | WCSA Magazine |  | Status bits | Positions | Commands
|  | HP-M300 Magazine |  | Status bits | Positions | Commands
11 | Robot
|  | ABB Tube Robot |  | Status bits | Positions | Commands
|  | ABB Preparation Robot |  | Status bits | Positions | Commands
|  | ABB Milling Circle Robot |  | Status bits | Positions | Commands
12 | Fusion Machines
|  | Rigaku XRF |  | Status bits | Positions | Commands
13 | Miscellaneous
|  | Timer |  | Status bits | Positions | Commands
|  | Slag Discharge Device |  | Status bits | Positions | Commands
|  | Transport System |  | Status bits | Positions | Commands

## Main
EQUIPMENT DETAILS
Sl.No | Unit |  | DETAILS
| Code | Name
STEEL INDUSTRY
1 | HR-ES / LS | Take over (Transfer Station) | Status bits | Positions | Commands
2 | HS-FF2000 | Milling Station | Status bits | Positions | Commands
3 | XRF | Spectrometer | Status bits | Positions | Commands
4 | HUST | Heating and Puching Machine | Status bits | Positions | Commands
5 | HS-CF | Milling & Cutting Station | Status bits | Positions | Commands
6 | OES | Optical Emission Spectrometer Analyser | Status bits | Positions | Commands
7 | HR-HSK | Sending & Receiving Station | Status bits | Positions | Commands
8 |  | Slag Discharge Device | Status bits | Positions | Commands
9 |  | Tube Robot | Status bits | Positions | Commands
10 |  | Preparation Robot | Status bits | Positions | Commands
11 | HP-CA | Automatic Crusher | Status bits | Positions | Commands
12 | HP-MA | Automatic Mill with Magnetic Separator for slag sample | Status bits | Positions | Commands
13 | HP-PA | Automatic Pelletizing Press for slag sample | Status bits | Positions | Commands
14 | CNS | Automatic CNS-Sample Magazine | Status bits | Positions | Commands
15 |  | Manual Gas Analyser | Status bits | Positions | Commands
16 |  | RadioActivity Measuring Device | Status bits | Positions | Commands
17 | Magazine | Input Cup Magazine | Status bits | Positions | Commands
18 | HS-F 3000 | 3-Axis Milling/Deburring Machine | Status bits | Positions | Commands
19 | Density | AccuPyc Density Meter | Status bits | Positions | Commands
20 | OES | OBLF OES | Status bits | Positions | Commands
21 | XRF | SuperQ XRF | Status bits | Positions | Commands
22 | Analyzer Database | Sympatec Analyzer Database | Status bits | Positions | Commands
MINING INDUSTRY
1 | WCSA | Scanner | Status bits | Positions | Commands
2 | WCSA | WCSA Magazine | Status bits | Positions | Commands
3 | WCSA | Lift | Status bits | Positions | Commands
4 | WCSA | Crusher | Status bits | Positions | Commands
5 | IRB-2600 | Robot - Milling Circle2 | Status bits | Positions | Commands
6 | IRB-2600 | Robot - Empty Cup | Status bits | Positions | Commands
7 | IRB-2600 | Robot - Buffer | Status bits | Positions | Commands
8 | HP-M1500 | Milling System | Status bits | Positions | Commands
9 | HP-SCD | Dosing System | Status bits | Positions | Commands
CEMENT INDUSTRY
1 | HR-BM | Sending Station | Status bits | Positions | Commands
2 | Timer | Timer | Status bits | Positions | Commands
3 | HR-LSP | Receiving Station | Status bits | Positions | Commands
4 | HP-M | Milling station | Status bits | Positions | Commands
5 | Transport System | Conveyor belt system for cup transport | Status bits | Positions | Commands
6 | Rigaku XRF | Fusion Machine | Status bits | Positions | Commands

## HR-ESLS
Name | Take Over | [5 empty] | BACK TO MAIN
Unit No | 121-122-123-124
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready |  |  | Register
2 | 1 | 2 | 2 | Remove Undefined Caps |  |  | start program
3 | 2 | 4 | 4 | Sample at Output |  |  | Lid opener
4 | 3 | 8 | 8 | Ready to receive Caps |  |  | cooling
5 | 4 | 10 | 16 | Ready to remove sample at output |  |  | drying
6 | 5 | 20 | 32 | Ready to Remove caps |  |  | Transport to ouput
7 | 8 | 100 | 256 | Syn is running |  |  | Ready for robot
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HR-ES/L 1 Take Over | 1 | Lid Opener | 1 | 10 | Start with turning
2 |  |  | 2 | Cooling | 2 | 11 | Start without turning
3 |  |  | 3 | Drying | 3 | 1001 | Start program 1 with turning
4 |  |  | 4 | Transport to Output | 4 | 1101 | Start program 1 without turning
5 |  |  | 5 | Ready for Robot | 5
6 |  |  | 6 | Caps. Not empty | 6
7 |  |  | 7 | Transport and air cooling | 7
8 |  |  | 50 | Start Empty caps | 50
9 |  |  | 51 | Start empty caps | 51
10 |  |  | 60 | Empty caps to HSK/L 1 | 60
11 |  |  | 61 | Empty caps to HSK/L 2 | 61
12 |  |  | 62 | Empty caps to HSK/L 3 | 62
13 |  |  | 63 | Empty caps to HSK/L 4 | 63
14 |  |  | 64 | Empty caps to HSK/L 5 | 64
15 |  |  | 65 | Empty caps to HSK/L 6 | 65
16 |  |  | 66 | Empty caps to HSK/L 7 | 66
17 |  |  | 67 | Empty caps to HSK/L 8 | 67
18 |  |  | 70 | Unknown Caps | 70
19 |  |  | 90 | Caps from robot | 90
20 |  |  | 91 | Empty caps to robot | 91

## HSFF-2000
Name | HSFF-2000 | [5 empty] | BACK TO MAIN
Unit No | 205
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready |  |  | Register
2 | 1 | 2 | 2 | Take Sample right allowed |  |  | Measure
3 | 2 | 4 | 4 | Bring Sample right allowed |  |  | Milling right/Milling left
4 | 3 | 8 | 8 | Robot Flap right is open |  |  | Ready for Output
5 | 4 | 10 | 16 | Manual Input is closed
6 | 5 | 20 | 32 | Manual Input is Open
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2056 | Special head inserted
11 | 12 | 1000 | 4096 | Robot flap right is closed
12 | 13 | 2000 | 8192 | Ready to open right flap after milling
13 | 14 | 4000 | 16384 | Deburr not possible
14 | 15 | 8000 | 32768 | Manual Input is locked
Unit No | 206 | HS-FF 2000_left flap
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Take sample left allowed
3 | 2 | 4 | 4 | Bring sample left allowed
4 | 3 | 8 | 8 | Robot flap left is open
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn is requested
7 | 10 | 400 | 1024 | Busy
8 | 12 | 1000 | 4096 | Robot flap left is closed
9 | 13 | 2000 | 8192 | Ready to open left flap after milling
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HSFF-2000 | 20 | Measure | 20 | 10 | Start Program
2 |  |  | 21 | Milling right | 21 | 11 | Open input flap
3 |  |  | 22 | Milling left | 22 | 12 | Open right robot-flap
4 |  |  | 23 | ready for output | 23 | 13 | Close right robot-flap
5 |  |  | 24 | sample too small | 24 | 14 | Lock input flap
6 |  |  | 25 | No Sample | 25 | 15 | Start Program, output left side
7 |  |  | 29 | Working | 29 | 22 | Open left robot-flap
8 |  |  | 30 | Start position Robot right | 30 | 23 | Close left robot-flap
9 |  |  | 31 | Startposition Manual | 31 | 24 | Move desk to right measuring device
10 |  |  | 32 | Startposition Robot left | 32 | 25 | Move desk to left measuring device
11 |  |  | 88 | Ready for out. left | 88
12 |  |  | 89 | Ready for out. right | 89
13 |  |  | 90 | Ready for robot | 90
Internal Unit 1 is defined in PM_no file but not specific positions and commands were mentioned

## HUST
Name | HUST | [5 empty] | BACK TO MAIN
Unit No | 195
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready |  |  | Sand blasting
2 | 2 | 4 | 4 | Disc sample in manual input |  |  | Heating
3 | 3 | 8 | 8 | Manual input is closed |  |  | Punching
4 | 4 | 10 | 16 | Tongue sample in manual input |  |  | Piece to turn table and waste
5 | 5 | 20 | 32 | CNS at storage
6 | 8 | 100 | 256 | Syn is running
7 | 9 | 200 | 512 | Syn is requested
8 | 10 | 400 | 1024 | Busy
Unit No | 196 | IN-OUT belt
1 | 0 | 1 | 1 | Ready
2 | 3 | 4 | 4 | Gripper open
3 | 4 | 8 | 8 | Gripper closed
4 | 8 | 10 | 16 | Syn is running
5 | 9 | 20 | 32 | Syn is requested
6 | 10 | 100 | 256 | Working
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HUST | 1 | Input | 10 | 3 | Reset
2 |  |  | 2 | CNS double thickness | 11 | 4 | Sync
3 |  |  | 3 | Sand blasting | 12 | 10 | Start Program double thickness sample from robot circle
4 |  |  | 4 | Heater double thickness | 13 | 11 | Start Program tongue sample from manual input to HORIBA
5 |  |  | 5 | Pin punch | 14 | 12 | Start Program disc sample from HS-CF Input
6 |  |  | 6 | Back at robot | 15 | 13 | Start Program discsample from manual input
7 |  |  | 7 | Sample too Long | 16
8 |  |  | 8 | Sample back at manual input | 17
9 |  |  | 9 | Heater Disc | 18
10 |  |  | 10 | CNS Disc | 19
11 |  |  | 11 | Waste box Disc | 20
12 |  |  | 50 | Manual input | 50
13 |  |  | 51 | Start from Robot | 51
14 |  |  | 52 | Start CNS Auto | 52
15 |  |  | 53 | Wait for CNS Auto | 53
16 |  |  | 54 | Start CNS Man | 54
17 |  |  | 55 | Wait for CNS Man | 55
18 |  |  | 95 | Start from Robot | 95
19 |  |  | 96 | Register tongue manual input | 96
20 |  |  | 97 | Register disc manual input | 97
21 |  |  | 98 | Register disc from HSCF | 98
22 | 2 | CNS-Magazine | 71 | Cup 1 | 71
23 |  |  | 72 | Cup 2 | 72
24 |  |  | 73 | Cup 3 | 73
25 |  |  | 74 | Cup 4 | 74
26 |  |  | 75 | Cup 5 | 75
27 |  |  | 76 | Cup 6 | 76
28 |  |  | 81 | Cup 1 | 81
29 |  |  | 82 | Cup 2 | 82
30 |  |  | 83 | Cup 3 | 83
31 |  |  | 84 | Cup 4 | 84
32 |  |  | 85 | Cup 5 | 85
33 |  |  | 86 | Cup 6 | 86

## HS-CF
Name | HS-CF | [5 empty] | BACK TO MAIN
Unit No | 180
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready |  |  | Cutting
2 | 1 | 2 | 2 | Milling not posssible |  |  | Cooling
3 | 2 | 4 | 4 | Breakdown |  |  | Drying
4 | 3 | 8 | 8 | Manual input is open |  |  | Millling
5 | 4 | 10 | 16 | Double thickness request |  |  | Deburring and output.
6 | 5 | 20 | 32 | Conical sample in manual input
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2056 | Bring calibration sample allowed
11 | 12 | 1000 | 4096 | Sample temperature too high
12 | 13 | 2000 | 8192 | Warning
13 | 14 | 4000 | 16384 | Standart sample in manual input
14 | 15 | 8000 | 32768 | Deburr not possible
Unit No | 181 | Disc Sample
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | ReCal request
3 | 4 | 10 | 16 | free
4 | 8 | 100 | 256 | free
Unit No | 182 | Automatic In-Output
1 | 0 | 1 | 1 | Ready
2 | 4 | 10 | 16 | Double thickness request
3 | 8 | 100 | 256 | Syn is running
4 | 10 | 400 | 1024 | Busy
5 | 11 | 800 | 2056 | Bring calibration sample allowed
6 | 12 | 1000 | 4096 | Sample temperature too high
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HS-CF1 | 20 | Measure | 20 | 9 | Start Program from robot for production sample
2 |  |  | 21 | Milling right | 21 | 10 | Start Program from robot for normal sample
3 |  |  | 22 | Milling left | 22 | 12 | Unlock input flap
4 |  |  | 23 | Ready for output | 23 | 13 | Start Program from manual for normal sample
5 |  |  | 24 | Sample too small | 24 | 14 | Start Program from manual for cut sample
6 |  |  | 25 | No sample | 25 | 19 | Start Program from robot for cut sample
7 |  |  | 26 | one Cutting | 26
8 |  |  | 27 | Watercooling | 27
9 |  |  | 28 | Deburrr | 28
10 |  |  | 29 | Working | 29
11 |  |  | 30 | Ready for robot | 30
12 |  |  | 31 | Startposition Manual | 31
13 |  |  | 33 | Two Cutting | 33
14 |  |  | 37 | Air cooling | 37
15 |  |  | 38 | In Transport | 38
16 |  |  | 90 | Ready for robot | 90
17 |  |  | 97 | Start from manual input pos | 97
18 |  |  | 98 | Start from rob pos | 98
19 | 1 | //Manual sample magazine, not used OR Disc sample | 93 | Disc sample in box | 93
20 | 2 | Automatic in-output | 31 | SampleOnBeltToHUST | 32
21 | 3 | CNS receiver | 91 | Chips in transport | 91
22 |  |  | 92 | Chips in cup | 92

## OES
Name | OES | [5 empty] | BACK TO MAIN
Unit No | 251
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready |  |  | Input the Sample
2 | 10 | 400 | 1024 | Reprep request/ask |  |  | Processing the Sample
3 | 11 | 800 | 2056 | Filing possible |  |  | Sample Finished
4 | 12 | 1000 | 4096 | SCT-session disabled
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | OES | 1 | Input | 10 | 10 | Register sample with Prog. = WS-program
2 |  |  | 2 | Input | 11 | 11 | Test reg. sample with Prog. 1
3 |  |  | 3 | Input | 12 | 12 | Test reg. sample with Prog. 2
4 |  |  | 4 | Finish | 20 | 40 | Delete Registration
5 |  |  | 5 | SCT-Finish | 21 | 50 | Sample prepared
6 |  |  | 6 | Finish | 22 | 51 | Sample prepared for SCT
7 |  |  | 7 | Output | 30
8 |  |  | 8 | Output | 31
9 |  |  | 9 | Output | 32
10 |  |  | 10 | Processing | 40
11 |  |  | 11 | Processing | 41
12 |  |  | 12 | Processing | 42
13 |  |  | 13 | Processing | 43
14 |  |  | 14 | Processing | 44
15 |  |  | 15 | Sample start | 50
16 |  |  | 16 | Sample on belt | 51
17 |  |  | 17 | Finish error | 24
18 |  |  | 18 | Output for re-prep | 18
19 |  |  | 19 | Output error | 34
20 |  |  | 50 | Starting | 55
21 |  |  | 51 | Measureing | 56
22 |  |  | 52 | Finished | 57
23 |  |  | 53 | Start SCT | 58
24 |  |  | 54 | Manual input | 59
25 |  |  | 55 | Manual input | 54
26 |  |  | 56 | Back from re-prep | 53
27 |  |  | 60 | Prepare starting | 60
28 |  |  | 80 | Analysing End Stat | 80
29 |  |  | 99 | Register... | 100
30 | 1 | Name not mentioned in ini file | 1 | Pos 1 | 61
31 |  |  | 2 | Pos 2 | 62
32 |  |  | 3 | Pos 3 | 63
33 |  |  | 4 | Pos 4 | 64
34 |  |  | 5 | Pos 5 | 65
35 |  |  | 6 | Pos 6 | 66
36 |  |  | 7 | Pos 7 | 67
37 |  |  | 8 | Pos 8 | 68
38 |  |  | 9 | Pos 9 | 69
39 |  |  | 10 | Pos 10 | 70

## HR-HSK L & B
Name | HR-HSK | HR-HSK/L 1 CM12, HR-HSK/B 1 CM12 |  |  |  |  | BACK TO MAIN
Unit No | 1
HR-HSK/L 1 CM12(Rec. Staion)
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Free2
3 | 3 | 8 | 8 | New caps. recived
4 | 4 | 10 | 16 | Input caps. OK
5 | 5 | 20 | 32 | station closed
6 | 8 | 100 | 256 | Syn is running
7 | 9 | 200 | 512 | Syn is requested
8 | 10 | 400 | 1024 | Operation
9 | 12 | 1000 | 4096 | free 12
10 | 13 | 2000 | 8192 | free 13
11 | 14 | 4000 | 16384 | Caps.corect in tube
| 41 | HR-HSK/B 1 CM12(Send. Station)
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Breakdown
3 | 3 | 8 | 8 | New capsule
4 | 4 | 10 | 16 | Caps. correct inside
5 | 5 | 20 | 32 | station closed
6 | 8 | 100 | 256 | Syn is running
7 | 9 | 200 | 512 | Syn is requested
8 | 10 | 400 | 1024 | Operation
9 | 11 | 1000 | 4096 | Station locked
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HR-HSK L | 11 | Capsule at RS | 11 | 1 | Stop
2 |  |  | 13 | Check capsule in lab | 13 | 2 | Reset
3 |  |  | 14 | Start in LAB Station | 14 | 3 | Switch to Automode
4 |  |  | 17 | Rmove Carrier | 17 | 4 | Switch to Man.mode
5 |  |  | 18 | Carrier removed | 18 | 5 | Start Sync.
6 |  |  | 53 | Send carrier back | 53 | 51 | Transp.to Lab.
7 |  |  | //90 | Empty capsule in LAB St. | 201 | 52 | Transp.to Send.st.
8 |  |  | //91 | Sample sticks in carrier? | 203 | 53 | Open tube for robot
9 |  |  | //51 | Take printout/Capsule | 204 | 54 | Open tube for operator
[6 empty] | //55 | Release Lab. Station
10 | 1 | HR-HSK B | 10 | Tube Transport | 10
11 |  |  | 12 | Capsule at SS | 12
12 |  |  | 15 | Start in plant | 15
13 |  |  | //99 | Registered | 200
14 |  |  | //98 | Round Iron with CNS (Start in plant) | 202

## HR-LSM
Name | HR-LSM | [5 empty] | BACK TO MAIN
Unit No | 1
HR-LSM 1 (Rec. Staion)
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Door closed
3 | 2 | 4 | 4 | Door locked
4 | 3 | 8 | 8 | Capsule arrived at lab
6 | 8 | 100 | 256 | Syn is running
7 | 9 | 200 | 512 | Syn is requested
8 | 10 | 400 | 1024 | Busy
9 | 11 | 800 | 2048 | Sensor 'Carrier arrived'
| 171 | Output station - Sample in Chute
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Breakdown
3 | 3 | 8 | 8 | Conical sample waiting
4 | 4 | 10 | 16 | Disc sample waiting
6 | 8 | 100 | 256 | Syn is running
7 | 9 | 200 | 512 | Syn is requested
8 | 10 | 400 | 1024 | Operation
9 | 11 | 800 | 2048 | Disc sample sucked during syn
9 | 12 | 1000 | 4096 | Conical sample in chute during syn
| 141 | Dummy unit 1
1 | 0 | 1 | 1 | Ready
| 151 | Dummy unit 2
| 161 | Dummy unit 3
| 121 | HR-HSK B (Sending Station)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Steel/pin Capsule
3 | 3 | 8 | 8 | New capsule insert Eason
4 | 4 | 10 | 16 | Capsule correct insert
6 | 5 | 20 | 32 | Station closed
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HR-LSM | 10 | On the way to lab | 10 | 1 | Stop
2 |  |  | 11 | Carrier arrived in lab | 11 | 2 | Reset
3 |  |  | 12 | Carrier back in sending station | 12 | 3 | Switch to Automode
4 |  |  | 13 | Empty capsule | 13 | 4 | Switch to Man.mode
5 |  |  | 16 | On the way to sending station | 16 | 5 | Start Sync.
6 |  |  | 17 | Sample in transport | 17 | 10 | Start sample output HS-SF Pos1
7 | 1 | Dummy 1 | 20 | HS-SF Transport Pos1 | 18 | 11 | Start sample output slag
8 |  |  | 21 | HS-SF Transport Pos2 | 19 | 12 | Start sample output pin to outlet tube
9 |  |  | 30 | Sample losed | 62 | 13 | Start sample output HS-SF Pos2
10 | 2 | Dummy 2 | 30 | Sample ready for transport | 20 | 14 | Start sample for pin sample to transport
11 |  |  | 23 | Pin Sample in cup | 21 | 15 | synchronisation steel output
12 |  |  | 50 | Waiting at outlet | 50 | 16 | synchronisation slag output
13 | 3 | Dummy 3 | 40 | ready for slag transport | 22 | 17 | steel transport back
14 |  |  | 41 | Moving to Slag | 23 | 18 | move steel from Pos2 to Pos1
15 | 4 | Output Station | 1 | Sample in chute | 25 | 19 | move steel from Pos1 to Pos2
16 |  |  | 2 | Conical sample available | 26 | 20 | Start sample output pin to cup
17 |  |  | 3 | Round sample available | 27 | 30 | Start sample to output to chute
18 |  |  | 90 | Sample in chute - Robot started | 90 | 31 | Select this LSM
19 |  |  | 91 | Unknown conical Sample in chute | 91 | 32 | Select other LSM
20 |  |  | 92 | Unknown disc Sample in sucker | 92 | 41 | Unlock front door
[6 empty] | 42 | Lock front door
[6 empty] | 51 | get capsule from station
[6 empty] | 52 | send capsule back to station
[6 empty] | 60 | Start conical sample to output
[6 empty] | 61 | Start round sample to output

## Slag Discharge Device
Name | STS Slag Discharge Device | [5 empty] | BACK TO MAIN
Unit No | 131
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Ready to put Cup
3 | 2 | 4 | 4 | Ready to remove cup
4 | 3 | 8 | 8 | Ready to put carrier
5 | 4 | 10 | 16 | Ready to remove carrier
6 | 5 | 20 | 32 | Undefine carrier available
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2056 | Undefine cup available
11 | 12 | 1000 | 4096 | Carrier cleaning necessary
12 | 13 | 2000 | 8192 | Funnel cleaning necessary
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | STS Slag Discharge Device | 1 | In Lid opener | 1 | 10 | Start Discharge sample
2 |  |  | 2 | Discharging | 2 | 11 | Start Carrier cleaning
3 |  |  | 3 | Ready for External | 3 | 12 | Start Funnel cleaning
4 |  |  | 10 | Start sample | 10
5 |  |  | 11 | Start carrier back | 11

## Tube Robot
Name | Tube Robot | [5 empty] | BACK TO MAIN
Unit No | 140
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Compressed air is ok
3 | 3 | 8 | 8 | Carrier present
4 | 4 | 10 | 16 | Lid present
5 | 5 | 20 | 32 | Steel carrier detected
Unit No | 141 | Robot 1 Sample lost
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | free_1
3 | 2 | 4 | 4 | free_2
4 | 3 | 8 | 8 | free_3
5 | 4 | 10 | 16 | free_4
6 | 5 | 20 | 32 | free_5
7 | 6 | 40 | 64 | free_6
8 | 7 | 80 | 128 | free_7
9 | 8 | 100 | 256 | free_8
10 | 9 | 200 | 512 | free_9
11 | 10 | 400 | 1024 | free_10
12 | 11 | 800 | 2056 | free_11
13 | 12 | 1000 | 4096 | free_12
14 | 13 | 2000 | 8192 | free_13
15 | 14 | 4000 | 16384 | free_14
16 | 15 | 8000 | 32768 | free_15
Unit No | 142 | Manual cup input 1
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | free_1
3 | 2 | 4 | 4 | Chute full cup
4 | 3 | 8 | 8 | Cup available
5 | 4 | 10 | 16 | free_4
6 | 5 | 20 | 32 | free_5
7 | 6 | 40 | 64 | free_6
8 | 7 | 80 | 128 | free_7
9 | 8 | 100 | 256 | free_8
10 | 9 | 200 | 512 | free_9
11 | 10 | 400 | 1024 | free_10
12 | 11 | 800 | 2056 | free_11
13 | 12 | 1000 | 4096 | free_12
14 | 13 | 2000 | 8192 | free_13
15 | 14 | 4000 | 16384 | free_14
16 | 15 | 8000 | 32768 | free_15
Unit No | 143 | Manual carrier input 1
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | free_1
3 | 2 | 4 | 4 | Chute full carrier
4 | 3 | 8 | 8 | Carrier available
5 | 4 | 10 | 16 | free_4
6 | 5 | 20 | 32 | free_5
7 | 6 | 40 | 64 | free_6
8 | 7 | 80 | 128 | free_7
9 | 8 | 100 | 256 | free_8
10 | 9 | 200 | 512 | free_9
11 | 10 | 400 | 1024 | free_10
12 | 11 | 800 | 2056 | free_11
13 | 12 | 1000 | 4096 | free_12
14 | 13 | 2000 | 8192 | free_13
15 | 14 | 4000 | 16384 | free_14
16 | 15 | 8000 | 32768 | free_15
Unit No | 144 | Cup magazine 1
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Cup Pos 1
3 | 2 | 4 | 4 | Cup Pos 2
4 | 3 | 8 | 8 | Cup Pos 3
5 | 4 | 10 | 16 | Cup Pos 4
6 | 5 | 20 | 32 | free_5
7 | 6 | 40 | 64 | free_6
8 | 7 | 80 | 128 | free_7
9 | 8 | 100 | 256 | free_8
10 | 9 | 200 | 512 | free_9
11 | 10 | 400 | 1024 | free_10
12 | 11 | 800 | 2056 | free_11
13 | 12 | 1000 | 4096 | free_12
14 | 13 | 2000 | 8192 | free_13
15 | 14 | 4000 | 16384 | free_14
16 | 15 | 8000 | 32768 | free_15
Unit No | 145 146 | Carrier output chute pos 1
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Error transfer
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Robot 1 | 1 | Pos1_Waste | 1
2 |  |  | //2 | Pos2_Carrier-Free | 2
3 |  |  | //3 | Pos3_Carrier-Free | 3
4 |  |  | 4 | Pos4_Carrier-Input | 4
5 |  |  | 5 | Pos5_Carrier-Output | 5
6 |  |  | 6 | Pos6_Carrier-PinOutput1 | 6
7 |  |  | 7 | Pos7_Carrier-PinOutput2 | 7
8 |  |  | 8 | Pos8_Carrier-CupMagazine | 8
9 |  |  | 9 | Pos9_Carrier-LidOpener | 9
10 |  |  | 10 | Pos10_Carrier-HSK1 | 10
11 |  |  | 11 | Pos11_Carrier-HSK2 | 11
12 |  |  | 12 | Pos12_Carrier-HSK3 | 12
13 |  |  | 13 | Pos13_Carrier-HSK4 | 13
14 |  |  | 14 | Pos14_Carrier-HSK5 | 14
15 |  |  | 15 | Pos15_Carrier-HSK6 | 15
16 |  |  | 16 | Pos16_Carrier-HSK7 | 16
17 |  |  | 17 | Pos17_Carrier-HSK8 | 17
18 |  |  | 18 | Pos18_Carrier-HSK9 | 18
19 |  |  | 19 | Pos19_Carrier-HSK10 | 19
20 |  |  | 20 | Pos20_Carrier_Steel-Transfer_1_1 | 20
21 |  |  | 21 | Pos21_Carrier_Conical-Transfer_1_3 | 21
22 |  |  | 22 | Pos22_Carrier-SlagSensor | 22
23 |  |  | 23 | Pos23_Carrier-SlagDischarger | 23
24 |  |  | 24 | Pos24_Carrier_Steel-Transfer_1_2 | 24
25 |  |  | 25 | Pos25_Carrier_Conical-Transfer_1_4 | 25
26 |  |  | //26 | Pos26_Carrier-Free | 26
27 |  |  | //27 | Pos27_Carrier-Free | 27
28 |  |  | //28 | Pos28_Carrier-Free | 28
29 |  |  | //29 | Pos29_Carrier-Free | 29
30 |  |  | //30 | Pos30_Carrier-Free | 30
31 |  |  | //31 | Pos31_Carrier-Free | 31
32 |  |  | 32 | Pos32_Cup-ManualInput | 32
33 |  |  | 33 | Pos33_Cup-ManualOutputChute | 33
34 |  |  | 34 | Pos34_Cup-MagazinePos1 | 34
35 |  |  | 35 | Pos35_Cup-MagazinePos2 | 35
36 |  |  | 36 | Pos36_Cup-MagazinePos3 | 36
37 |  |  | 37 | Pos37_Cup-MagazinePos4 | 37
38 |  |  | 38 | Pos38_Cup-BeltPos | 38
39 |  |  | 39 | Pos39_Cup-SlagDischarger | 39
40 |  |  | 40 | Pos40_Cup-Free | 40
41 |  |  | 41 | Pos41_Cup-Free | 41
42 |  |  | 42 | Pos42_Cup-Free | 42
43 |  |  | 43 | Pos43_Cup-Free | 43
44 |  |  | 44 | Pos44_Cup-Free | 44
45 |  |  | 45 | Pos45_Cup-Free | 45
46 |  |  | 46 | Pos46_Cup-Free | 46
47 |  |  | 47 | Pos47_Cup-Free | 47
48 |  |  | 48 | Pos48_Cup-Free | 48
49 |  |  | 49 | Pos49_Cup-Free | 49
50 |  |  | 90 | carrier is OK | 211
51 |  |  | 91 | carrier is not OK | 212
52 |  |  | 99 | Robot | 99
54
55 | 1 | Robot 1 Sample lost | 1 | Lost at Pos1_Waste | 101
56 |  |  | //2 | Lost at Pos2_Carrier-Free | 102
57 |  |  | //3 | Lost at Pos3_Carrier-Free | 103
58 |  |  | 4 | Lost at Pos4_Carrier-Input | 104
59 |  |  | 5 | Pos5_Carrier-Output | 105
60 |  |  | 6 | Pos6_Carrier-PinOutput1 | 106
61 |  |  | 7 | Pos7_Carrier-PinOutput2 | 107
62 |  |  | 8 | Pos8_Carrier-CupMagazine | 108
63 |  |  | 9 | Pos9_Carrier-LidOpener | 109
64 |  |  | 10 | Pos10_Carrier-HSK1 | 110
65 |  |  | 11 | Pos11_Carrier-HSK2 | 111
66 |  |  | 12 | Pos12_Carrier-HSK3 | 112
67 |  |  | 13 | Pos13_Carrier-HSK4 | 113
68 |  |  | 14 | Pos14_Carrier-HSK5 | 114
69 |  |  | 15 | Pos15_Carrier-HSK6 | 115
70 |  |  | 16 | Pos16_Carrier-HSK7 | 116
71 |  |  | 17 | Pos17_Carrier-HSK8 | 117
72 |  |  | 18 | Pos18_Carrier-HSK9 | 118
73 |  |  | 19 | Pos19_Carrier-HSK10 | 119
74 |  |  | 20 | Pos20_Carrier_Steel-Transfer1 | 120
75 |  |  | 21 | Pos21_Carrier_Conical-Transfer2 | 121
76 |  |  | 22 | Lost at Pos22_Carrier-SlagDischarger | 122
77 |  |  | //23 | Lost at Pos23_Carrier-Free | 123
78 |  |  | //24 | Lost at Pos24_Carrier-Free | 124
79 |  |  | //25 | Lost at Pos25_Carrier-Free | 125
80 |  |  | //26 | Lost at Pos26_Carrier-Free | 126
81 |  |  | //27 | Lost at Pos27_Carrier-Free | 127
82 |  |  | //28 | Lost at Pos28_Carrier-Free | 128
83 |  |  | //29 | Lost at Pos29_Carrier-Free | 129
84 |  |  | //30 | Lost at Pos30_Carrier-Free | 130
85 |  |  | //31 | Lost at Pos31_Carrier-Free | 131
86 |  |  | 32 | Lost at Pos32_Cup-ManualInput | 132
87 |  |  | 33 | Lost at Pos33_Cup-ManualOutputChute | 133
88 |  |  | 34 | Lost at Pos34_Cup-MagazinePos1 | 134
89 |  |  | 35 | Lost at Pos35_Cup-MagazinePos2 | 135
90 |  |  | 36 | Lost at Pos36_Cup-MagazinePos3 | 136
91 |  |  | 37 | Lost at Pos37_Cup-MagazinePos4 | 137
92 |  |  | 38 | Lost at Pos38_Cup-BeltPos | 138
93 |  |  | 39 | Lost at Pos39_Cup-SlagDischarger | 139
94 |  |  | 40 | Lost at Pos40_Cup-Free | 140
95 |  |  | 41 | Lost at Pos41_Cup-Free | 141
96 |  |  | 42 | Lost at Pos42_Cup-Free | 142
97 |  |  | 43 | Lost at Pos43_Cup-Free | 143
98 |  |  | 44 | Lost at Pos44_Cup-Free | 144
99 |  |  | 45 | Lost at Pos45_Cup-Free | 145
100 |  |  | 46 | Lost at Pos46_Cup-Free | 146
101 |  |  | 47 | Lost at Pos47_Cup-Free | 147
102 |  |  | 48 | Lost at Pos48_Cup-Free | 148
103 |  |  | 49 | Lost at Pos49_Cup-Free | 149
104
105 | 2 |  | 1 | Cup input | 200
106 |  |  | 2 | Outputchute | 201
107 |  |  | //3 | Manual input cup
108
109 | 3 |  | 1 | Carrier Input | 202
110 |  |  | 2 | Car. chute pos.4 | 203
111 |  |  | 3 | Car. chute pos.3 | 204
112
113 | 4 |  | 1 | Cup on park pos 1 | 205
114 |  |  | 2 | Cup on park pos 2 | 206
115 |  |  | 3 | Cup on park pos 3 | 207
116 |  |  | 4 | Cup on park pos 4 | 208
117
118 | 5 |  | 1 | Car. chute pos.2 | 209
119
120 | 6 |  | 1 | Car. chute pos.1 | 210

## Preparation Robot
Name | Tube Robot | 4 nos of unit |  |  |  |  | BACK TO MAIN
Unit No | 160 | Robot 3
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Compressed air is ok
3 | 3 | 8 | 8 | Water is OK
Unit No | 161 | Robot Sample lost
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | free_1
3 | 2 | 4 | 4 | free_2
4 | 3 | 8 | 8 | free_3
5 | 4 | 10 | 16 | free_4
6 | 5 | 20 | 32 | free_5
7 | 6 | 40 | 64 | free_6
8 | 7 | 80 | 128 | free_7
9 | 8 | 100 | 256 | free_8
10 | 9 | 200 | 512 | free_9
11 | 10 | 400 | 1024 | free_10
12 | 11 | 800 | 2056 | free_11
13 | 12 | 1000 | 4096 | free_12
14 | 13 | 2000 | 8192 | free_13
15 | 14 | 4000 | 16384 | free_14
16 | 15 | 8000 | 32768 | free_15
Unit No | 162 | Manual input R3
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Input closed
3 | 3 | 8 | 8 | Input locked
4 | 4 | 10 | 16 | Sample at pos 1
5 | 5 | 20 | 32 | Sample at pos 2
6 | 6 | 40 | 64 | Chute full
Unit No | 163 | Sample magazine R3
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Sample at pos 1
3 | 2 | 4 | 4 | Sample at pos 2
4 | 3 | 8 | 8 | Sample at pos 3
5 | 4 | 10 | 16 | Sample at pos 4
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Robot 1 | 1 | Pos1_Waste | 1
2 |  |  | 8 | Pos8_CoolingUnit | 8
3 |  |  | 9 | Pos9_TransferToOtherCircle | 9
4 |  |  | 10 | Pos10_TransferToOtherCircle-Turned | 10
5 |  |  | 11 | Pos11_HSFF_Round | 11
6 |  |  | 12 | Pos12_HSFF_DoubleThickness | 12
7 |  |  | 13 | Pos13_MagazinePos1 | 13
8 |  |  | 14 | Pos14_MagazinePos2 | 14
9 |  |  | 15 | Pos15_MagazinePos3 | 15
10 |  |  | 16 | Pos16_MagazinePos4 | 16
11 |  |  | //17 | Pos17 Mag at OES_1 | 17
12 |  |  | //18 | Pos18 Mag at OES_2 | 18
13 |  |  | //19 | Pos19 Mag at OES_3 | 19
14 |  |  | 20 | Pos20_OE1_Input | 20
15 |  |  | 21 | Pos21_OE1_InputTurned | 21
16 |  |  | 22 | Pos22_OE1_Output | 22
17 |  |  | 23 | Pos23_OE1_OutputTurned | 23
18 |  |  | 24 | Pos24_XRF1_Input | 24
19 |  |  | 25 | Pos25_XRF1_InputTurned | 25
20 |  |  | 26 | Pos26_XRF1_Output | 26
21 |  |  | 27 | Pos27_XRF1_OutputTurned | 27
22 |  |  | 32 | Pos32_HSCFInputKonicalSample | 32
23 |  |  | 33 | Pos33_HSCFIn-Output | 33
24 |  |  | //34 | Pos34_Steel_HUST | 34
25 |  |  | 35 | Pos35_Steel_TO_1_oval | 35
26 |  |  | 36 | Pos36_Steel_TO_2_oval | 36
27 |  |  | 37 | Pos37_TO_3_conical | 37
28 |  |  | 38 | Pos38_TO_4_conical | 38
29 |  |  | 39 | Pos39_Steel_TO_1_oval_turned | 39
30 |  |  | 40 | Pos40_Steel_TO_2_oval_turned | 40
31 |  |  | 41 | Pos41_In-/Output Belt HUST 1 | 41
32 |  |  | 42 | Pos42_magazine belt | 42
33 |  |  | 43 | Pos37_TO_3_Double thickness | 43
34 |  |  | 44 | Pos44_TO_4_Double thickness | 44
35 |  |  | 99 | Robot | 99
36
37 | 1 | Robot Sample lost | 1 | Lost at Pos1_Waste | 101
38 |  |  | //2 | Lost at Pos2_Carrier-Free | 100
39 |  |  | //3 | Lost at Pos3_Carrier-Free | 100
40 |  |  | 4 | Lost atPos4_Sample-ManualInputTongue | 104
41 |  |  | 5 | Lost atPos5_Sample-ManualOutputChute | 105
42 |  |  | 6 | Lost atPos6_Sample-ManualInputRound | 106
43 |  |  | 7 | Lost atPos7_Sample-ManualOutputRing | 107
44 |  |  | 8 | Lost atPos8_Sample-CoolingUnit | 108
45 |  |  | 9 | Lost atPos9_CSample-TransferUnitToOtherCircle | 109
46 |  |  | 10 | Lost atPos10_Sample-TransferUnitToOtherCircle-Turned | 110
47 |  |  | 11 | Lost atPos11_Sample-HSFF_Round | 111
48 |  |  | 12 | Lost atPos12_Sample-HSFF_DoubleThickness | 112
49 |  |  | 13 | Lost atPos13_Sample-MagazinePos1 | 113
50 |  |  | 14 | Lost atPos14_Sample-MagazinePos2 | 114
51 |  |  | 15 | Lost atPos15_Sample-MagazinePos3 | 115
52 |  |  | 16 | Lost atPos16_Sample-MagazinePos4 | 116
53 |  |  | 17 | Lost atPos17 Mag at OES_1 | 117
54 |  |  | 18 | Lost atPos18 Mag at OES_2 | 118
55 |  |  | 19 | Lost atPos19 Mag at OES_3 | 119
56 |  |  | 20 | Lost atPos20_Steel_OE1_Input | 120
57 |  |  | 21 | Lost atPos21_Steel_OE1_InputTurned | 121
58 |  |  | 22 | Lost atPos22_Steel_OE1_Output | 122
59 |  |  | 23 | Lost atPos23_Steel_OE1_OutputTurned | 123
60 |  |  | 24 | Lost atPos24_Steel_XRF1_Input | 124
61 |  |  | 25 | Lost atPos25_Steel_XRF1_InputTurned | 125
62 |  |  | 26 | Lost atPos26_Steel_XRF1_Output | 126
63 |  |  | 27 | Lost atPos27_Steel_XRF1_OutputTurned | 127
64 |  |  | 28 | Lost atPos28_Steel_OE3_Input | 128
65 |  |  | 29 | Lost atPos29_Steel_OE3_InputTurned | 129
66 |  |  | 30 | Lost atPos30_Steel_OE3_Output | 130
67 |  |  | 31 | Lost atPos31_Steel_OE3_OutputTurned | 131
68 |  |  | 32 | Lost atPos32_Steel_HSCFInputKonicalSample | 132
69 |  |  | 33 | Lost atPos33_Steel_HSCFIn-Output | 133
70 |  |  | 34 | Lost atPos34_Steel_HUST | 134
71 |  |  | 35 | LostatPos35_Steel_Transferpos1 | 135
72 |  |  | 36 | LostatPos36_Steel_Transferpos2 | 136
73 |  |  | 37 | Lost atPos37_Steel_Transferpos5 | 137
74 |  |  | 38 | Lost atPos38_Steel_Transferpos6 | 138
75 |  |  | 39 | Lost atPos39_Iron_MagazineBelt | 139
76 |  |  | 40 | Lost atPos40_Iron_XRF1 | 140
77 |  |  | 41 | Lost atPos41_In-/Output Belt HUST 1 | 141
78 |  |  | 42 | Lost atPos42_magazine belt | 142
79 |  |  | 43 | Lost atPos43_TO_3_Double thickness | 143
80 |  |  | 44 | Lost atPos44_TO_4_Double thickness | 144
81 |  |  | 45 | Lost atPos45_HSFFInputDoubleThickness | 145
82 |  |  | 46 | Lost atPos46_HSFFOutputDoubleThickness | 146
83 |  |  | 47 | Lost atPos47_Iron_RingMagazinePos2 | 147
84 |  |  | 48 | Lost atPos48_Iron_RingMagazinePos3 | 148
85 |  |  | 49 | Lost atPos49_Iron_RingMagazinePos4 | 149
86 |  |  | 50 | Lost atPos50_Iron_HB3000-1 | 150
87 |  |  | 51 | Lost atPos51_Iron_HB3000-2 | 151
88 |  |  | 52 | Lost atPos52_Iron_Transferposition1 | 152
89 |  |  | 53 | Lost atPos53_Iron_Transferposition1-Turned | 153
90 |  |  | 54 | Lost atPos54_Iron_Transferposition2 | 154
91 |  |  | 55 | Lost atPos55_Iron_Transferposition2-Turned | 155
92 |  |  | 56 | Lost atPos56_Iron_Free | 156
93 |  |  | 57 | Lost atPos57_Iron_Free | 157
94 |  |  | 58 | Lost atPos58_Iron_Free | 158
95 |  |  | 59 | Lost atPos59_Iron_Free | 159
96
97 | 2 | Manual input R3 | 1 | Sample input
98 |  |  | 2 | Outputchute | 200
99 |  |  | //3 | Manual input cup | 201
100 |  |  | 5 | Pos5_ManualOutputChute | 5
101 |  |  | 6 | Pos6_ManualInputRound | 6
102 |  |  | 7 | Pos7_ManualOutputConcial | 7
106
107 | 3 | Sample magazine R3 | 1 | Sample on park pos 1 | 202
108 |  |  | 2 | Sample on park pos 2 | 203
109 |  |  | 3 | Sample on park pos 3 | 204
110 |  |  | 4 | Sample on park pos 4 | 205
111 |  |  | 5 | Calibration Mag 1 | 210
112 |  |  | 6 | Calibration Mag 2 | 211
113 |  |  | 7 | Calibration Mag 3 | 212
114
115 | 4 |  | 1 | Ring on park pos 1 | 206
116 |  |  | 2 | Ring on park pos 2 | 207
117 |  |  | 3 | Ring on park pos 3 | 208
118 |  |  | 4 | Ring on park pos 4 | 209

## Robot_Commands
ABB Robot
| Default | Source Position. Defined in the Robot | Source Magazine Position | Destination Position. Defined in the Robot | Destination Magazine Position
Commands (Value - 0102000300) | 01 | 02 | 00 | 03 | 00
BACK TO MAIN
|  | Source Position. Defined in the Robot | Source Magazine Position | Destination Position. Defined in the Robot | Destination Magazine Position
Commands (Value - 0102000300) | 01 | 02 | 00 | 03 | 00

## Chute Magazine
Name | Chute Magazine | [5 empty] | BACK TO MAIN
Unit No | 208 | Name=Magazine
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Sample at output XRF
3 | 3 | 8 | 8 | Setup mode
4 | 4 | 10 | 16 | All chutes full
5 | 8 | 100 | 256 | Syn in operation
6 | 9 | 200 | 512 | Syn is requested
7 | 10 | 400 | 1024 | In operation
8 | 12 | 1000 | 4096 | Sample at printer
9 | 13 | 2000 | 8192 | Sample at output OES1
10 | 14 | 4000 | 16384 | Sample at output OES3
11 | 15 | 8000 | 32768 | Sample at output OES4
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 1 | Magazine | 1 | Chute 1 | 1 | 10010101 | Sample from OES 1 to chute 1
2 |  |  | 2 | Chute 2 | 2 | 10010201 | Sample from OES 1 to chute 2
3 |  |  | 3 | Chute 3 | 3 | 10010301 | Sample from OES 1 to chute 3
4 |  |  | 4 | Chute 4 | 4 | 10010401 | Sample from OES 1 to chute 4
5 |  |  | 5 | Chute 5 | 5 | 10010501 | Sample from OES 1 to chute 5
6 |  |  | 6 | Chute 6 | 6 | 10010601 | Sample from OES 1 to chute 6
7 |  |  | 7 | Chute 7 | 7 | 10010701 | Sample from OES 1 to chute 7
8 |  |  | 8 | Chute 8 | 8 | 10010801 | Sample from OES 1 to chute 8
9 |  |  | 9 | Chute 9 | 9 | 10010901 | Sample from OES 1 to chute 9
10 |  |  | 10 | Chute 10 | 10 | 10011001 | Sample from OES 1 to chute 10
11 |  |  | 11 | Chute 11 | 11 | 10011101 | Sample from OES 1 to chute 11
12 |  |  | 12 | Chute 12 | 12 | 10011201 | Sample from OES 1 to chute 12
13 |  |  | 13 | Chute 13 | 13 | 10011301 | Sample from OES 1 to chute 13
14 |  |  | 14 | Chute 14 | 14 | 10011401 | Sample from OES 1 to chute 14
15 |  |  | 15 | Chute 15 | 15 | 10011501 | Sample from OES 1 to chute 15
16 |  |  | 16 | Chute 16 | 16 | 10011601 | Sample from OES 1 to chute 16
17 |  |  | 17 | On Transport belt | 17 | 10050101 | Sample from start pisition of belt to chute 1
18 |  |  | 19 | Printing | 19 | 10050201 | Sample from start pisition of belt to chute 2
19 |  |  | 20 | Printing finished | 20 | 10050301 | Sample from start pisition of belt to chute 3
20 |  |  | 21 | waste box end of belt | 21 | 10050401 | Sample from start pisition of belt to chute 4
[6 empty] | 10050501 | Sample from start pisition of belt to chute 5
[6 empty] | 10050601 | Sample from start pisition of belt to chute 6
[6 empty] | 10050701 | Sample from start pisition of belt to chute 7
[6 empty] | 10050801 | Sample from start pisition of belt to chute 8
[6 empty] | 10050901 | Sample from start pisition of belt to chute 9
[6 empty] | 10051001 | Sample from start pisition of belt to chute 10
[6 empty] | 10051101 | Sample from start pisition of belt to chute 11
[6 empty] | 10051201 | Sample from start pisition of belt to chute 12
[6 empty] | 10051301 | Sample from start pisition of belt to chute 13
[6 empty] | 10051401 | Sample from start pisition of belt to chute 14
[6 empty] | 10051501 | Sample from start pisition of belt to chute 15
[6 empty] | 10051601 | Sample from start pisition of belt to chute 16
[6 empty] | 10030101 | Sample from OES 3 to chute 1
[6 empty] | 10030201 | Sample from OES 3 to chute 2
[6 empty] | 10030301 | Sample from OES 3 to chute 3
[6 empty] | 10030401 | Sample from OES 3 to chute 4
[6 empty] | 10030501 | Sample from OES 3 to chute 5
[6 empty] | 10030601 | Sample from OES 3 to chute 6
[6 empty] | 10030701 | Sample from OES 3 to chute 7
[6 empty] | 10030801 | Sample from OES 3 to chute 8
[6 empty] | 10030901 | Sample from OES 3 to chute 9
[6 empty] | 10031001 | Sample from OES 3 to chute 10
[6 empty] | 10031101 | Sample from OES 3 to chute 11
[6 empty] | 10031201 | Sample from OES 3 to chute 12
[6 empty] | 10031301 | Sample from OES 3 to chute 13
[6 empty] | 10031401 | Sample from OES 3 to chute 14
[6 empty] | 10031501 | Sample from OES 3 to chute 15
[6 empty] | 10031601 | Sample from OES 3 to chute 16
[6 empty] | 10040101 | Sample from OES 4 to chute 1
[6 empty] | 10040201 | Sample from OES 4 to chute 2
[6 empty] | 10040301 | Sample from OES 4 to chute 3
[6 empty] | 10040401 | Sample from OES 4 to chute 4
[6 empty] | 10040501 | Sample from OES 4 to chute 5
[6 empty] | 10040601 | Sample from OES 4 to chute 6
[6 empty] | 10040701 | Sample from OES 4 to chute 7
[6 empty] | 10040801 | Sample from OES 4 to chute 8
[6 empty] | 10040901 | Sample from OES 4 to chute 9
[6 empty] | 10041001 | Sample from OES 4 to chute 10
[6 empty] | 10041101 | Sample from OES 4 to chute 11
[6 empty] | 10041201 | Sample from OES 4 to chute 12
[6 empty] | 10041301 | Sample from OES 4 to chute 13
[6 empty] | 10041401 | Sample from OES 4 to chute 14
[6 empty] | 10041501 | Sample from OES 4 to chute 15
[6 empty] | 10041601 | Sample from OES 4 to chute 16

## Cup Magazine
Name | Input Cup Magazine | [5 empty] | BACK TO MAIN
Unit No | 32 | Name=Input cup Magazine
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
| 0 | 1 | 1 | Ready
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 1 | Input Cup Magazine | 1 | Next Sample | 1 | 36 | Insert a sample
2 |  |  | 2 | Register | 2 | 37 | Remove a sample
3 |  |  | 3 | Starting | 3 | 38 | Insert sample (internal)
[6 empty] | 39 | choise a free position (internal)

## HP-CA
Name | HP-CA 1 | [5 empty] | BACK TO MAIN
Unit No | 214
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
3 | 2 | 4 | 4 | Bring cup allowed
5 | 4 | 10 | 16 | Cup inside machine
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
11 | 12 | 1000 | 4096 | Cup in push input
12 | 13 | 2000 | 8192 | Push input is closed
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HP-CA 1 | 1 | Output | 16 | 10 | Sample from left to input
2 |  |  | 2 | In Crusher | 17 | 11 | Sample from right to middel
3 |  |  | 3 | On belt at lift | 18 | 12 | Cup from pushinput to input
4 |  |  | 4 | On belt left | 19 | 13 | Cup from input to left
5 |  |  | 5 | On belt right | 20 | 14 | Cup from input to right
6 |  |  | 6 | Push Input | 21 | 15 | Cup from input to pushinput
7 |  |  | //7 | Backup Sample filled in Cup | 25 | 16 | Cup-lift up after crushing
8 |  |  | //8 | Transport Cup at Low Position | 26 | 17 | Start crushing from belt
9 |  |  | //9 | Main Sample ready | 27 | 18 | Switch to auto mode
10 |  |  | 22 | On Belt left | 22 | 19 | Switch to manual mode
11 |  |  | //23 | Cup in magazin | 23 | 20 | Start synchronisation
12 |  |  | //24 | No cup in magazin | 24 | 21 | Start cleaning
14 |  |  | 90 | Wait for start | 90 | 22 | Reset breakdown
15 |  |  | 91 | Sample at input belt | 91 | 23 | Stop machine
16 |  |  | 92 | Transport Cup back | 92 | 24 | Search for next cup
17 |  |  | 93 | Wait for filling | 93 | 27 | Start crushing from LK
18 |  |  | 98 | Registered at Push Input | 98 | 30 | Sample Cup from Input to Bottom
[6 empty] | 31 | Transport Cup to Output
[6 empty] | 32 | Start dummy Sample up to belt
[6 empty] | 33 | Sample through machine right to left
[6 empty] | 34 | Sample through machine left to right

## HP-MA
Name | HP-MA 1 | [5 empty] | BACK TO MAIN
Unit No | 215
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Allowed to take a sample (robot)
3 | 2 | 4 | 4 | Allowed to bring a sample (robot)
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn reqested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2056 | Breakdown
11 | 12 | 1000 | 4096 | Warning
12 | 13 | 2000 | 8192 | Cup in Push Input
13 | 14 | 4000 | 16384 | Push Inp. Closed
14 | 15 | 8000 | 32768 | Cup on belt at magazine
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HP-MA1 | 1 | Output | 16 | 10 | Sample from left to input
2 |  |  | 2 | Grinding | 17 | 11 | Sample from right to input
3 |  |  | 3 | On belt at lift | 18 | 12 | Cup from pushinput to input
4 |  |  | 4 | On belt | 19 | 13 | Cup from input to left
5 |  |  | 5 | Dosing | 20 | 14 | Cup from input to right
6 |  |  | 6 | Manual input | 21 | 15 | Cup from input to pushinput
7 |  |  | 7 | Cup back | 22 | 16 | Cup-lift up after grinding
8 |  |  | 8 | Magnetic separator | 24 | 17 | Start grinding
9 |  |  | 9 | Push Input | 51 | 18 | Switch to auto mode
10 |  |  | 10 | On belt right to press | 25 | 19 | Switch to manual mode
11 |  |  | 30 | Belt Start | 30 | 20 | Start synchronisation
12 |  |  | 50 | Start Mill Manu Input | 50 | 21 | Start cleaning
13 |  |  | 51 | Sample registred | 52 | 22 | Reset breakdown
14 |  |  | 90 | Wait for start | 90 | 23 | Stop machine
15 |  |  | 91 | Empty Cup back | 91 | 31 | Make space in cup magazine
16 |  |  | 94 | Wait for empty Cup in Input | 94 | 32 | Search cup magazine
17 |  |  | 98 | Registered at Push Input | 98 | 33 | Cup through MA2 Left(CA)->Right(PA)
[6 empty] | 34 | Cup through MA2 Right(PA)->Left(CA)
[6 empty] | 35 | Start program from left side on belt

## HP-PA
Name | HP-PA | [5 empty] | BACK TO MAIN
Unit No | 216 | HP-PA
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Cup at end of belt
3 | 2 | 4 | 4 | Ready to bring
4 | 3 | 8 | 8 | Input belt free
5 | 4 | 10 | 16 | Cup in manual input
6 | 5 | 20 | 32 | Cup input closed
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn reqested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2056 | Breakdown
11 | 12 | 1000 | 4096 | Warning
12 | 13 | 2000 | 8192 | Ring cleaning idle
13 | 14 | 4000 | 16384 | Ready to take a cup (robot)
14 | 15 | 8000 | 32768 | Ready to bring a cup (robot)
Unit No | 217 | Name=Ring cleaning
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Park3
3 | 2 | 4 | 4 | Park4
4 | 3 | 8 | 8 | Ring cleaning idle (output allowed)
5 | 4 | 10 | 16 | Stackmagazine refill active
6 | 8 | 100 | 256 | Syn is running
7 | 10 | 400 | 1024 | Busy
Unit No | 218 | Name=HP-PA 1 output belt
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Ready to bring a sample
3 | 2 | 4 | 4 | Ready to take a sample
4 | 3 | 8 | 8 | No Ring at end
5 | 4 | 10 | 16 | Ring in press on output belt
6 | 8 | 100 | 256 | Syn is running
7 | 9 | 200 | 512 | Syn necessary
8 | 10 | 400 | 1024 | Busy
9 | 11 | 800 | 2048 | Ring at end of belt- XRF1
10 | 12 | 1000 | 4096 | Ring at end of belt- XRF2
11 | 13 | 2000 | 8192 | Ring at end of belt- XRF
12 | 14 | 4000 | 16384 | Ring output magazine is full
13 | 15 | 8000 | 32768 | Ring on return belt
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HP-PA | 2 | Ready in press tool | 21 | 10 | Start press from belt
2 |  |  | 3 | On output belt | 22 | 11 | Start press from input
3 |  |  | 5 | In press | 24 | 12 | Cup from belt to push input
4 |  |  | 6 | Push Input | 25 | 13 | Cup from belt back to mill
5 |  |  | 7 | Cup on belt | 26 | 14 | Start sample output (from presstool to belt)
6 |  |  | 8 | Middle of belt | 27 | 15 | Sample from press to XRF
7 |  |  | 10 | Cleaning in operation | 29 | 16 | Start ringcleaning
8 |  |  | 11 | Transport belt | 30 | 17 | Start ring back from return-chute
9 |  |  | 12 | Cupcleaning | 31 | 18 | Start cup cleaning
10 |  |  | 13 | Empty cup back | 32 | 19 | Cup from middle of inputbelt -> Pushinput
11 |  |  | 16 | Lack of material | 35 | 20 | Switch to auto mode
12 |  |  | 17 | Cup return position | 50 | 21 | Switch to manual mode
13 |  |  | //20 | Return Cup | 39 | 22 | Start synchronisation
14 |  |  | //21 | Wait for start | 40 | 23 | Reset breakdown
15 |  |  | //22 | Empty Cup back | 41 | 24 | Stop machine
16 |  |  | 23 | Empty Cup to cup slider | 42 | 25 | Ring Press > Magazine(extern)
17 |  |  | 24 | Empty Cup in cup slider | 43 | 26 | Ring Magazine(extern) > Analyser
18 |  |  | 50 | Cup return position | 50 | 27 | Cup right to cup slider
19 |  |  | 51 | Start Cup at Press | 51 | 28 | Lock autostart ring stack magazine
20 |  |  | 81 | In input | 81 | 29 | Unlock autostart ring stack magazine
21 |  |  | 98 | Registered at Push Input | 200 | 30 | Empty Cup back to magazine via cup slider
22 | 1 | Ring Cleaning | 1 | Ring cleaning | 20 | 38 | Start ringcleaning from XRF-2
23 |  |  | 4 | Waste bin | 36 | 39 | Sample from press to XRF-2
24 | 2 | HP-PA Output belt
25 |  |  | 9 | End of belt XRF | 90
26 |  |  | 14 | Sample on transport to XRF | 33
27 |  |  | 15 | Ring in chute | 34
28 |  |  | 17 | Ring return chute | 38
29 |  |  | 18 | Ring in turn table Magazine | 37
30 |  |  | 23 | End of belt XRF-1
31 |  |  | 24 | End of belt XRF-2
32 |  |  | 81 | Sample in analyser | 82
33 |  |  | 90 | Sample at analyser | 83

## XRF
Name | XRF | [5 empty] | BACK TO MAIN
Unit No | 250
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
2 | 10 | 400 | 1024 | Reprep request/ask
3 | 11 | 800 | 2056 | Filing possible
4 | 12 | 1000 | 4096 | SCT-session disabled
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | XRF-1 | 1 | Input | 10 | 10 | Register sample with Prog. = WS-program
2 |  |  | 2 | Input | 11 | 11 | Test reg. sample with Prog. 1
3 |  |  | 3 | Input | 12 | 12 | Test reg. sample with Prog. 2
4 |  |  | 4 | Finish | 20 | 40 | Delete Registration
5 |  |  | 5 | SCT-Finish | 21 | 50 | Sample prepared
6 |  |  | 6 | Finish | 22 | 51 | Sample prepared for SCT
7 |  |  | 7 | Output | 30
8 |  |  | 8 | Output | 31
9 |  |  | 9 | Output | 32
10 |  |  | 10 | Processing | 40
11 |  |  | 11 | Processing | 41
12 |  |  | 12 | Processing | 42
13 |  |  | 13 | Processing | 43
14 |  |  | 14 | Processing | 44
15 |  |  | 15 | Sample start | 50
16 |  |  | 16 | Sample on belt | 51
17 |  |  | 17 | Finish error | 24
18 |  |  | 18 | Output for re-prep | 18
19 |  |  | 19 | Output error | 34
20 |  |  | 50 | Starting | 55
21 |  |  | 51 | Measureing | 56
22 |  |  | 52 | Finished | 57
23 |  |  | 53 | Start SCT | 58
24 |  |  | 54 | Manual input | 59
25 |  |  | 55 | Manual input | 54
26 |  |  | 56 | Back from re-prep | 53
27 |  |  | 60 | Prepare starting | 60
28 |  |  | 80 | Analysing End Stat | 80
29 |  |  | 99 | Register | 99
30 | 1 | Name not Mentioned in pm_no_250.INI file nor declared but position written | 1 | Pos 1 | 61
31 |  |  | 2 | Pos 2 | 62
32 |  |  | 3 | Pos 3 | 63
33 |  |  | 4 | Pos 4 | 64
34 |  |  | 5 | Pos 5 | 65
35 |  |  | 6 | Pos 6 | 66
36 |  |  | 7 | Pos 7 | 67
37 |  |  | 8 | Pos 8 | 68
38 |  |  | 9 | Pos 9 | 69
39 |  |  | 10 | Pos 10 | 70

## CNS-Magazine
Name | Turntable CNS Magazine | [5 empty] | BACK TO MAIN
Unit No | 240
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
4 | 3 | 8 | 8 | Setup mode
7 | 8 | 100 | 256 | Syn in operation
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Turntable CNS Magazine | 10 | Transport | 10 | 1 | Stop
2 |  |  | 11 | Arrived | 11 | 2 | Reset
3 |  |  | //12 | Cup transport belt | 12 | 3 | Setup mode
4 |  |  | //13 | Cup at belt end | 13 | 4 | Automatic
5 |  |  | //14 | Cup at take over position | 14 | 5 | Start syn.
6 |  |  | //15 | Cup at lift CA, steel | 15 | 10 | Start transport from HUST 1
7 |  |  | //16 | Cup at lift CA, iron | 16 | 11 | Start transport from HUST 2

## CNS-Analyzer
Name | Leco CNS Analyzer | [5 empty] | BACK TO MAIN
Unit No | 260
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
2 | 10 | 400 | 1024 | Sample Detected
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Turntable CNS Magazine | 1 | Measuring | 1 | 10 | Register sample with Prog. = WS-program
2 |  |  | 2 | Completing SOP | 2 | 11 | Register sample with Prog. 1
3 |  |  | 50 | Start | 50 | 12 | Register sample with Prog. 2
4 |  |  | 51 | Starting | 51 | 13 | Register sample with Prog. 3
5 |  |  | 52 | Done | 52 | 14 | Register sample with Prog. 4
6 | [5 empty] | 15 | Register sample with Prog. 5
7 | [5 empty] | 16 | Register sample with Prog. 6
[6 empty] | 17 | Register sample with Prog. 7
[6 empty] | 18 | Register sample with Prog. 8
[6 empty] | 19 | Register sample with Prog. 9
[6 empty] | 102 | Analyze sample

## Manual Gas analyser
Name | Manual Leco Gas Analyser | [5 empty] | BACK TO MAIN
Unit No | 261
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
9 | 10 | 400 | 1024 | Sample Detected
10
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Manual Leco Gas Analyser | 1 | Measuring | 1 | 10 | Register sample with Prog. = WS-program
2 |  |  | 2 | Completing SOP | 2 | 11 | Register sample with Prog. 1
3 |  |  | 50 | Start | 50 | 12 | Register sample with Prog. 2
4 |  |  | 51 | Starting | 51 | 13 | Register sample with Prog. 3
5 |  |  | 52 | Done | 52 | 14 | Register sample with Prog. 4
[6 empty] | 15 | Register sample with Prog. 5
[6 empty] | 16 | Register sample with Prog. 6
[6 empty] | 17 | Register sample with Prog. 7
[6 empty] | 18 | Register sample with Prog. 8
[6 empty] | 19 | Register sample with Prog. 9
[6 empty] | 102 | Analyze sample

## RMS device
Name | RMS | [5 empty] | BACK TO MAIN
Unit No | 271
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec) |  |  |  | FUNCTION
1 | 0 | 1 | 1 | Ready
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | RMS | 1 | Sample introduced: No sample | 1 | 11 | Start production sample measurement with Shape1
2 |  |  | 2 | Sample introduced: Buffer full | 2 | 12 | Start production sample measurement with Shape2
3 |  |  | 3 | Result: Sample clean | 3 | 13 | Start production sample measurement with Shape3
4 |  |  | 4 | Result: Radioactivity alarm detected | 4 | 21 | Start Control Sample measurement
5 |  |  | 5 | CS-Result: System on error | 5 | 22 | Start Control Sample measurement with Shape2
6 |  |  | 6 | CS-Result: System operational | 6 | 23 | Start Control Sample measurement with Shape3
7 |  |  | 10 | At RMS (register) | 10

## Barcode Scanner
Name | Barcode Scanner | [5 empty] | BACK TO MAIN
Unit No | 18
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Barcode Scanner | 1 | Reading barcode | 1 | 10 | Start reader (laser on)
2 |  |  | 2 | Read ok | 2 | 11 | Stop reader (laser off)
3 |  |  | 3 | Read failed | 3 | 12 | Lock laser
4 |  |  | 10 | Chk Priority | 10 | 13 | Unlock laser
5 |  |  | 11 | Chk Duplicate | 11 | 14 | Test-Modus 1 (Reading speed / number of successful decodes)
6 |  |  | 15 | Get Cup Dupl/1 | 15 | 15 | Test-Modus 2 (Number of decodes the barcode is visible during)
7 |  |  | 16 | Get Cup Normal/1 | 16 | 16 | Quit testmode
8 |  |  | 25 | Get Cup Dupl/2 | 25 | 17 | Turn the OK output on
9 |  |  | 26 | Get Cup Normal/2 | 26 | 18 | Turn the NG output on
10 |  |  | 35 | Get Cup Dupl/3 | 35 | 19 | Turn the OK/NG outputs off
11 |  |  | 36 | Get Cup Normal/3 | 36 | 20 | Delete transmission buffer
12 |  |  | 50 | Start reader | 50
13 |  |  | 52 | Read ok | 52
14 |  |  | 53 | Read failed | 53
15 |  |  | 55 | Request Vials ... | 55
16 |  |  | 56 | Request Vials HP-SCD/1 | 56
17 |  |  | 57 | Request Vials HP-SCD/2 | 57
18 |  |  | 58 | Request Vials HP-SCD/3 | 58
19 |  |  | 59 | Request Vials HP-SCD OK | 59
20 |  |  | 69 | Request2 Vials HP-SCD OK | 69

## Scanner
Name | Scanner | [5 empty] | BACK TO MAIN
Unit No | 55
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2
3 | 2 | 4 | 4
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Scanner_WCSA_1 | 1 | Waiting | 1
2 |  | Scanner_WCSA_1 | 2 | Processing | 2
3 |  | Scanner_WCSA_1 | 3 | Pos | 3

## WCSA_Mag
Name | WCSA Magazine | [5 empty] | BACK TO MAIN
Unit No | 63
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Manual Mode
3 | 4 | 10 | 16 | Insertation Inactive
4 | 5 | 20 | 32 | Scan of sample possible
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn is requested
7 | 10 | 400 | 1024 | Busy
8 | 11 | 800 | 2048 | Magazine is full
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | WCSA_MAG 1 | 1 | Prepare Bucket | 1 | 1 | Stop
2 |  |  | 2 | No Bucket | 2 | 10 | Insert Sample
3 |  |  | 3 | Empty bucket to heavy | 3 | 11 | Reorganize
4 |  |  | 4 | Wait for sample | 4 | 12 | Bucket to WCSA
5 |  |  | 5 | Check weight | 5 | 13 | Empty Bucket to Buffer
6 |  |  | 6 | Weight Ok | 6 | 2 | Reset
7 |  |  | 7 | Weight NOT ok | 7 | 3 | Manual Mode
8 |  |  | 8 | Bucket in magazine | 8 | 4 | Auto Mode
9 |  |  | 9 | Bucket Out during Syn | 9 | 5 | Syn
10 |  |  | 11 | Bucket to Lift | 11
11 |  |  | 12 | Bucket in Lift lower position | 12
12 |  |  | 13 | Bucket in Transport | 13
13 |  |  | 80 | Unknow bucket return belt | 231
14 | 1 | WCSA_MAG 1 Internal | 1 | Starting .... | 201
15 |  |  | 10 | Pos 1 | 210
16 |  |  | 11 | Pos 2 | 211
17 |  |  | 12 | Pos 3 | 212
18 |  |  | 13 | Pos 4 | 213
19 |  |  | 14 | Pos 5 | 214
20 |  |  | 15 | Pos 6 | 215
21 |  |  | 16 | Pos 7 | 216
22 |  |  | 17 | Pos 8 | 217
23 |  |  | 18 | Pos 9 | 218
24 |  |  | 19 | Pos 10 | 219
25 |  |  | 20 | Pos 11 | 220
26 |  |  | 21 | Pos 12 | 221
27 |  |  | 22 | Pos 13 | 222
28 |  |  | 23 | Pos 14 | 223
29 |  |  | 24 | Pos 15 | 224
30 |  |  | 25 | Pos 16 | 225
31 |  |  | 26 | Pos 17 | 226
32 |  |  | 27 | Pos 18 | 227
33 |  |  | 28 | Pos 19 | 228
34 |  |  | 29 | Pos 20 | 229
35 |  |  | 30 | Pos 21 | 230

## WCSA_Lift
Name | WCSA_Lift | [5 empty] | BACK TO MAIN
Unit No | 65
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Nil
3 | 2 | 4 | 4 | Nil
4 | 3 | 8 | 8 | Manual Mode
5 | 4 | 10 | 16 | Lift is in Upper position
6 | 5 | 20 | 32 | Lift is in Lower position
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2048 | Bucket in lower lift position
11 | 12 | 1000 | 4096 | Ready to take from Sky Transport
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | WCSA_1 lift | 1 | Bucket at bottom position | 1 | 1 | Stop
2 |  | WCSA_1 lift | 2 | Bucket in transport | 2 | 10 | Lift bucket to Crusher Input
3 |  | WCSA_1 lift | 3 | Bucket ready to dump | 3 | 11 | Deplete material into Crusher
4 |  | WCSA_1 lift | 4 | Bucket empty | 4 | 12 | Prepare for sky loading
5 |  | WCSA_1 lift | 5 | Inserting | 5 | 13 | From sky takeover to crusher input
6 |  | WCSA_1 lift | 98 | Bucket from sky transport | 199 | 2 | Reset
7 |  | WCSA_1 lift | 99 | Bucket at Bottom | 200 | 3 | Manual Mode
[6 empty] | 4 | Auto Mode
[6 empty] | 5 | Syn
[6 empty] | 12 | Prepare for sky loading

## WCSA
Name | WCSA | [5 empty] | BACK TO MAIN
Unit No | 6
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Cup from splitter outlet 3 @ robot position
3 | 2 | 4 | 4 | Sample has been inserted
4 | 3 | 8 | 8 | Manual mode
5 | 4 | 10 | 16 | Cup from Splitter outlet 1 at robot positiion
6 | 5 | 20 | 32 | Nil
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2048 | Ready to insert sample
11 | 12 | 1000 | 4096 | Material in Cup under outlet 1
12 | 13 | 2000 | 8192 | Material in Cup under outlet 2
13 | 14 | 4000 | 16384 | Material in Cup under outlet 3
14 | 15 | 8000 | 32768 | Cup from Splitter oultet 2 at Robot posistinon
Unit No | 61 | Transport from Chamber1
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Cup under outlet 3 filled ?????
3 | 4 | 10 | 16 | Cup from splitter outlet 1 at robot pos
4 | 8 | 100 | 256 | Syn is running
5 | 9 | 200 | 512 | Syn is requested
6 | 10 | 400 | 1024 | Busy
7 | 11 | 800 | 2048 | Cup under outlet 1 filled
8 | 13 | 2000 | 8192 | Transport at robot
9 | 14 | 4000 | 16384 | Transport at machine
10 | 15 | 8000 | 32768 | Cup for outlet 1 available at Machine
Unit No | 62 | Transport from Chamber2
1 | 0 | 1 | 1 | Ready
2 | 4 | 10 | 16 | Cup from splitter outlet 2 at robot pos
3 | 8 | 100 | 256 | Syn is running
4 | 9 | 200 | 512 | Syn is requested
5 | 10 | 400 | 1024 | Busy
6 | 11 | 800 | 2048 | Cup under outlet 2 filled
7 | 13 | 2000 | 8192 | Transport at robot
8 | 14 | 4000 | 16384 | Transport at machine
9 | 15 | 8000 | 32768 | Cup for outlet 2 available at Machine
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | WCSA_1 | 1 | Inserting | 1 | 1 | Stop
2 |  |  | 2 | Crushing | 2 | 10 | Start with Cup 1,2,3
3 |  |  | 3 | Cup is Full | 3 | 11 | Start with cup1,-,3
4 |  |  | 4 | Cup in Transport | 4 | 12 | Start with Cup 1,2,--
5 |  |  | 5 | Sample at Robot | 5 | 13 | Start with Cup 1,-,-
6 |  |  | 6 | Sample from chamber 3 at robot | 6 | 2 | Reset
7 | 1 | Transport from chamber 1 | 1 | Filled | 240 | 3 | Manual Mode
8 |  | Transport from chamber 1 | 2 | Ready | 242 | 4 | Auto Mode
9 |  | Transport from chamber 1 | 90 | Returning | 250 | 5 | Syn
10 |  | Transport from chamber 1 | 99 | Unknown | 252
11 | 2 | Transport from chamber 2 | 1 | Filled | 241
12 |  | Transport from chamber 2 | 2 | Ready | 243
13 |  | Transport from chamber 2 | 90 | Returning | 251
14 |  | Transport from chamber 2 | 99 | Unknown | 253

## Combo-machine
Name | HP-BT (Combo) | [5 empty] | BACK TO MAIN
Unit No | 8
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Setup mode
3 | 4 | 10 | 16 | Transport at HP-BT
4 | 5 | 20 | 32 | Transport at robot
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn is requested
7 | 10 | 400 | 1024 | Busy
8 | 14 | 4000 | 16384 | Cup on transport
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HP-BT(Combo) | 1 | Bucket received | 1 | 1 | Stop
2 |  |  | 2 | Bucket in transport | 2 | 2 | Reset
3 |  |  | 3 | Bucket drop pos. | 3 | 3 | Manual Mode
4 |  |  | 4 | Crushing/Splitting | 4 | 4 | Auto mode
5 |  |  | 5 | Sample at robot | 5 | 5 | Syn
[6 empty] | 13 | Start splitting, Rest to waste
[6 empty] | 14 | Start splitting, Rest to outputbag

## Robot Milling Circle2
Name | Robot_ Milling Circle2 |  |  | Node |  |  | BACK TO MAIN
Unit No | 9 |  |  | 9
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 8 | 100 | 256 | Syn is running
3 | 9 | 200 | 512 | Syn is requested
4 | 10 | 400 | 1024 | Busy
5 | 11 | 800 | 2048
6 | 12 | 1000 | 4096
Unit No | 91 | Sample lost
1 | 0 | 1 | 1 | Ready
Unit No | 92 | Input-Output
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Input closed
Unit No | 93 | Cleaning
1 | 0 | 1 | 1 | Ready
Unit No | 94 | Buffer
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | CupMag 1 present
3 | 2 | 4 | 4 | CupMag 2 present
4 | 3 | 8 | 8 | CupMag 3 present
5 | 4 | 10 | 16 | CupMag 4 present
6 | 5 | 20 | 32 | CupMag 5 present
7 | 6 | 40 | 64 | CupMag 6 present
8 | 7 | 80 | 128 | CupMag 7 present
9 | 8 | 100 | 256 | CupMag 8 present
10 | 9 | 200 | 512 | CupMag 9 present
11 | 10 | 400 | 1024 | CupMag 10 present
12 | 11 | 800 | 2048 | CupMag 11 present
13 | 12 | 1000 | 4096 | CupMag 12 present
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Robot | 1 | Pos1_Waste | 1
2 |  | Robot | 2 | Pos2_ManInput | 2
3 |  | Robot | 3 | Pos3_ManOutputChute | 3
4 |  | Robot | 4 | Pos4_Cleaning | 4
5 |  | Robot | 6 | Pos6_HP-SCD | 6
6 |  | Robot | 7 | Pos7_HP-M1500_1 | 7
7 |  | Robot | 8 | Pos7_HP-M1500_2 | 8
8 |  | Robot | 9 | Pos7_HP-M1500_3 | 9
9 |  | Robot | 10 | Pos7_HP-M1500_4 | 10
10 |  | Robot | 11 | Pos11_HP-WCSA_1_Cup_1 | 11
11 |  | Robot | 12 | Pos11_HP-WCSA_1_Cup_2 | 12
12 |  | Robot | 13 | Pos11_HP-WCSA_2_Cup_1 | 13
13 |  | Robot | 14 | Pos11_HP-WCSA_2_Cup_2 | 14
14 |  | Robot | 15 | Pos11_HP-WCSA_3_Cup_1 | 15
15 |  | Robot | 16 | Pos11_HP-WCSA_3_Cup_2 | 16
16 |  | Robot | 99 | Roboter | 99
17 | 1 | Sample lost | 1 | Pos1_Waste | 101
18 |  |  | 2 | Pos2_ManInput | 102
19 |  |  | 3 | Pos3_ManOutputChute | 103
20 |  |  | 4 | Pos4_Cleaning | 104
21 |  |  | 6 | Pos6_HP-SCD | 106
22 |  |  | 7 | Pos7_HP-M1500_1 | 107
23 |  |  | 8 | Pos8_HP-M1500_2 | 108
24 |  |  | 9 | Pos9_HP-M1500_3 | 109
25 |  |  | 10 | Pos10_HP-M1500_4 | 110
26 |  |  | 11 | Pos11_HP-WCSA_1_Cup_1 | 111
27 |  |  | 12 | Pos12_HP-WCSA_1_Cup_2 | 112
28 |  |  | 13 | Pos13_HP-WCSA_2_Cup_1 | 113
29 |  |  | 14 | Pos14_HP-WCSA_2_Cup_2 | 114
30 |  |  | 15 | Pos15_HP-WCSA_3_Cup_1 | 115
31 |  |  | 16 | Pos16_HP-WCSA_3_Cup_2 | 116
32 | 2 | Input-Output | 1 | Start regular | 200
33 |  |  | 2 | Start Fusion only | 201
34 |  |  | 3 | Start empty cup | 202
35 |  |  | 4 | Output chute | 203
36 | 3 | Cleaning
37 | 4 | Buffer | 1 | Pos21_Cup_Park1 | 21
38 |  |  | 2 | Pos22_Cup_Park2 | 22
39 |  |  | 3 | Pos23_Cup_Park3 | 23
40 |  |  | 4 | Pos24_Cup_Park4 | 24
41 |  |  | 5 | Pos25_Cup_Park5 | 25
42 |  |  | 6 | Pos26_Cup_Park6 | 26
43 |  |  | 7 | Pos27_Cup_Park7 | 27
44 |  |  | 8 | Pos28_Cup_Park8 | 28
45 |  |  | 9 | Pos29_Cup_Park9 | 29
46 |  |  | 10 | Pos30_Cup_Park10 | 30
47 |  |  | 11 | Pos31_Cup_Park11 | 31
48 |  |  | 12 | Pos32_Cup_Park12 | 32

## HP-M1500
Name | HP-M1500 |  |  | Node |  |  | BACK TO MAIN
Unit No | 2 |  |  | 2
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Take Allowed
3 | 2 | 4 | 4 | Bring allowed
4 | 8 | 100 | 256 | Syn. Running
5 | 9 | 200 | 512 | Syn.Necessary
6 | 10 | 400 | 1024 | Busy
7 | 11 | 800 | 2048
8 | 12 | 1000 | 4096
9 | 15 | 8000 | 32768 | Input Allowed
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HP-M1500 1 | 1 | Input Cycle | 20 | 10 | Start Grinding
2 |  | HP-M1500 1 | 2 | Dosing | 21 | 11 | Start Cleaning
3 |  | HP-M1500 1 | 3 | Grinding | 22 | 20 | Switch to Automatic
4 |  | HP-M1500 1 | 4 | Output | 23 | 21 | Switch to Manual
5 |  | HP-M1500 1 | 5 | Ready/1 | 24 | 22 | Start Syn
6 |  | HP-M1500 1 | 6 | Cup at Output | 25 | 23 | Reset
7 |  | HP-M1500 1 | 7 | Ready weight not ok | 26 | 24 | Stop
8 |  | HP-M1500 1 | 8 | Balance | 27
9 |  | HP-M1500 1 | 10 | Wait to Start | 250
10 |  | HP-M1500 1 | 52 | Remove cup | 52
11 |  | HP-M1500 1 | 90 | Ready/2 | 251

## HP-SCD
Name | HP-SCD |  |  | Node |  |  | BACK TO MAIN
Unit No | 10 |  |  | 10
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Take big cup allowed
3 | 2 | 4 | 4 | Bring big cup allowed
4 | 3 | 8 | 8 | Setup Mode
5 | 8 | 100 | 256 | Syn. Running
6 | 9 | 200 | 512 | Syn.is requested
7 | 10 | 400 | 1024 | Busy
8 | 14 | 4000 | 16384 | Cleaning Sampler requested
9 | 15 | 8000 | 32768 | Little cup on exhaust
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HP-SCD | 1 | Starting | 1 | 1 | Stop
2 |  | HP-SCD | 2 | Sampling done | 2 | 10 | Fill Sampler
3 |  | HP-SCD | 4 | Cup-transport | 4 | 13 | Cleaning
4 |  | HP-SCD | 5 | Cup at Transport | 5 | 2 | Reset
5 |  | HP-SCD | 6 | Cup at exhaust dosing | 6 | 3 | Setup Mode
6 |  | HP-SCD | 7 | Sampling | 7 | 4 | Automatic
7 |  | HP-SCD | 50 | Start | 50 | 5 | Start Syn
8 |  | HP-SCD | 51 | Start | 51
9 |  | HP-SCD | 52 | Remove cup | 52
10 |  | HP-SCD | 53 | Waiting | 53

## HP-M300 Magazine
Name | HP-M300 Magazine | [5 empty] | BACK TO MAIN
Unit No | 15 | 300 pos. magazine
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Magazine 2 available
3 | 2 | 4 | 4 | Magazine 3 available
4 | 3 | 8 | 8 | Setup mode
5 | 4 | 10 | 16 | Breakdown magazine
6 | 5 | 20 | 32 | Magazine 1 available
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | In operation
10 | 11 | 800 | 2048 | Ready to scan
Unit No | 151 | Output belt (Belt Magazine <-----> HAG-6 HF)
11 | 0 | 1 | 1 | Ready
12 | 1 | 2 | 2 | Ready to lift cup
13 | 3 | 8 | 8 | Setup mode
14 | 4 | 10 | 16 | Breakdown magazine
15 | //5 | 20 | 32 | spare
16 | 8 | 100 | 256 | Syn.Running
17 | 9 | 200 | 512 | Syn is requested
18 | 10 | 400 | 1024 | In operation
19 | 11 | 800 | 2048 | Cup at top of SCD lift
Unit No | 152 | input belt (HP-SCD <-->Magazine)
20 | 0 | 1 | 1 | Ready
21 | 3 | 8 | 8 | Setup mode
22 | 4 | 10 | 16 | Breakdown magazine
23 | 8 | 100 | 256 | Syn.Running
24 | 9 | 200 | 512 | Syn is requested
25 | 10 | 400 | 1024 | In operation
26 | //11 | 800 | 2048 | spare
27 | 12 | 1000 | 4096 | Full cup at HAG
28 | 13 | 2000 | 8192 | Holder for empty cup at HAG
29 | 14 | 4000 | 16384 | Empty cup at HAG
Unit No | 153 | Cleaning
30 | 0 | 1 | 1 | Ready
31 | 3 | 3 | 8 | Setup mode
32 | 8 | 100 | 256 | Syn.Running
33 | 10 | 400 | 1024 | In operation
34 | 12 | 1000 | 4096 | Vial present
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 15 | 300 pos Magazine | 1 | XY-Gripper | 1 | 1 | Stop
2 |  | Robot | 2 | Vial return HPM | 2 | 2 | Reset
3 |  | Robot | 4 | At Cleaning | 4 | 3 | Setup mode
4 |  | Robot | 5 | Scanner | 5 | 4 | Automatic
5 |  | Robot | 6 | Scan ready | 6 | 5 | Start syn.
6 |  | Robot | 7 | Mag1 scanned | 7 | 10 | Clean cup and scan barcode
7 |  | Robot | 8 | Mag2 scanned | 8 | 11 | Cup from barcode scanner back to magazine
8 |  | Robot | 16 | Mag3 Scanned | 16 | 13 | Cup from HPM300 to HAG
9 |  | Robot | 18 | No vial at magazine | 18 | 14 | Return cup from HAG to HPM300
10 |  | Robot | 19 | HAG Vial stored | 19 | 15 | Move cups into HP-SCD
11 |  |  | 80 | Request Vials/1 | 200 | 16 | Move cups out of HP-SCD
12 |  |  | 81 | Request Vials/2 | 201 | 17 | Move cups from HPM300 -> HP-SCD
13 |  |  | 82 | Request Vials/3 | 201 | 18 | Move cups from HPM300 -> HP-SCD with Cleaning
14 |  |  | 83 | Request Vials/4 | 201 | 19 | Move cups from HPM300 -> HP-SCD with Cleaning and Scanner
15 |  |  | 84 | Request Vials/5 | 201 | 20 | Return cups from HP-SCD to HPM300
16 |  |  | 85 | Request 1 Vial | 202 | 22 | Clean vial
17 |  |  | 86 | Request 2 Vial | 203 | 23 | Move cleaning->scanner
18 |  |  | 87 | Request 3 Vial | 204 | 24 | Move cup from HP-SCD to Magazine
19 |  | Robot | 90 | Register Mag1 | 210
20 |  | Robot | 91 | Register Mag2 | 211
21 |  | Robot | 92 | Register Mag3 | 212
22 |  | Robot | 93 | Register Mag4 | 213
23 |  | Robot | 94 | Register Done | 214
24 | 151 | input belt (HP-SCD <-->Magazine) | 1 | SCD-Transport | 9
25 |  |  | 2 | At SCD | 10
26 |  |  | 3 | Under HPM | 3
27 |  |  | 10 | Check/1 | 200
28 |  |  | 11 | Wait/1 | 201
29 |  |  | 12 | Wait for HPM/1 | 202
30 |  |  | 20 | Check/2 | 203
31 |  |  | 21 | Wait/2 | 204
32 |  |  | 22 | Wait for HPM/2 | 205
33 |  |  | 30 | Check/3 | 206
34 |  |  | 31 | Wait/3 | 207
35 |  |  | 33 | Wait for HPM/3 | 208
36 |  |  | 40 | Done | 209
37 | 152 | Output belt (Belt Magazine <--> HAG-6 HF) | 13 | HAG-Transport | 13
38 |  |  | 14 | At HAG | 14
39 |  |  | 15 | Transport from HAG | 15
40 | 153 | Cleaning | 21 | Cleaning .. | 11
41 |  |  | 22 | Cleaning OK | 12

## Tray Magazine
Name | Buffer Tray Magazine | [5 empty] | BACK TO MAIN
Unit No | 50,51,52,53
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Item Present
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Tray Magazine | 1 | Import Tray | 1 | 10 | Insert a sample
2 |  |  | 2 | Register | 2 | 11 | Remove a sample
3 |  |  | 3 | Starting ... | 3 | 20 | Insert sample (internal)
4 |  |  | 11 | Next sample | 11 | 30 | choise a free position (internal)
5 |  |  | 12 | Next vial | 12 | 50 | Check Magazine position is free
6 |  |  | 13 | Next empty vial | 13 | 41 | Mag sorting order=Priority
7 |  |  | 14 | Next successfull sample | 14 | 42 | Mag sorting order=Register order
8 |  |  | 15 | Register vial to carrier/1 | 15 | 43 | Mag sorting order=First in First out
9 |  |  | 16 | Register vial to carrier/2 | 16
10 |  |  | 17 | Register vial to carrier/3 | 17
11 |  |  | 20 | Perpare | 20
12 |  |  | 21 | Tray | 21
13 |  |  | 22 | Register | 22

## HR-BM
Name | HR-BM | Sending station |  |  |  |  | BACK TO MAIN
Unit No | 35
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
4 | 3 | 8 | 8 | Not enough material inside mixer
5 | 4 | 10 | 16 | Carrier not empty
6 | 5 | 20 | 32 | Carrier Present
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn is requested
9 | 10 | 400 | 1024 | Busy
10 | 13 | 2000 | 8192 | Warning
| 36 | Sampler 1(HR-SN)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Ext. sampler in operation
3 | 2 | 4 | 4 | Ext sampler breakdown
4 | 3 | 8 | 8 | Buffer full
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn is requested
7 | 10 | 400 | 1024 | Busy
8 | 15 | 8000 | 32768 | Ext. sampler ready for operation
| 37 | Sampler 2
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Buffer Full
3 | 8 | 100 | 256 | Syn is running
4 | 9 | 200 | 512 | Syn is requested
5 | 10 | 400 | 1024 | Busy
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HR-BM | 6 | Mixer | 26 | 11 | Start sampling to funnel 1
2 |  |  | 7 | Dosing Tube | 27 | 12 | Start dosing from funnel 1
3 |  |  | 8 | Filled in carrier | 28 | 13 | Cancel sampling 1 (clean buffer and mixer)
4 |  |  | 9 | Sample to waste | 29 | 14 | Start spot sampling 1
5 |  |  | 10 | Capsule ready to send | 30 | 15 | Stop sampling 1
6 |  |  | 15 | Sample in bottle | 35 | //16 | Start sample to bottle from funnel 1
7 |  |  | 16 | Sample in bottle 1 | 36 | 21 | Start sampling to funnel 2
8 |  |  | 17 | Sample in bottle 2 | 37 | 22 | Start dosing from funnel 2
9 |  |  | 20 | carrier waste material | 38 | 23 | Cancel sampling 2 (clean buffer and mixer)
10 |  |  | 60 | Next Carrier | 60 | 24 | Start spot sampling 2
11 |  |  | 51 | Sampling done | 51 | 25 | Stop sampling 2
12 |  |  | 52 | Sampling cancelled | 52 | //26 | Start sample to bottle from funnel 2
13 |  |  | 90 | Clean Mixer | 53 | 50 | Start filling capsule
14 |  |  | 91 | Clean wait | 54 | 51 | Start cleaning
15 | [5 empty] | 52 | Stop spot sample
16 | 1 | Sampler(HR-SN) | 1 | Sampling Funnel1 | 22
17 |  |  | 11 | Sampling finished | 31
18 |  |  | 61 | Start spot | 61
19 |  |  | 90 | Canceled, wait for cleaning | 90
20
21 | 2 | Sampler 2 | 1 | Sampling Funnel2 | 23
22 |  |  | 11 | Sampling finished | 32
23 |  |  | 61 | Start spot | 62
24

## Timer
Name | Timer Mill | [5 empty] | BACK TO MAIN
Unit No | 83
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Timer Mill | 1 | Start sampling | 20 | 10 | Start sample Taking normal sample
2 |  |  | 11 | Finnished. Normal sample | 11 | 11 | Start sample Taking spot sample
3 |  |  | 12 | Finnished. Spot sample | 12 | 30 | Stop sample taking normal sample
4 |  |  | 31 | Cancel. Normal sample | 31 | 31 | Stop sample taking spot sample
5 |  |  | 32 | Cancel. Spot sample | 32
6 |  |  | 51 | Finnished. Normal sample | 51
7 |  |  | 52 | Finnished. Spot sample | 52

## HR-LSP
Name | HR-LSP | Receiving station |  |  |  |  | BACK TO MAIN
Unit No | 7 | HP-LSP 1
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Take cup allowed
3 | 2 | 4 | 4 | Put cup allowed
4 | 3 | 8 | 8 | cleaning requested
5 | 4 | 10 | 16 | Cleaning allowed
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn reqested
9 | 10 | 400 | 1024 | Busy
12 | 13 | 2000 | 8192 | Dedusting is running
13 | 14 | 4000 | 16384 | Ready to give out cup
14 | 15 | 8000 | 32768 | Carrier is in lab
| 71 | HR-LSP Output
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Big cup available
3 | 3 | 8 | 8 | Take allowed
4 | 4 | 10 | 16 | Put allowed
5 | 5 | 20 | 32 | Dosing allowed
6 | 8 | 100 | 256 | Syn is running
7 | 10 | 400 | 1024 | Busy
8 | 13 | 2000 | 8192 | Cup under output or dosing
| 72 | Tube-1
1 | 0 | 1 | 1 | Ready
2 | 1 | 2 | 2 | Carrier present
3 | 2 | 4 | 4 | Diverter 1 Pos. A
4 | 3 | 8 | 8 | Diverter 1 Pos. B
5 | 4 | 10 | //16 | Ready to send
6 | 5 | 20 | //32 | Carrier Present
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn reqested
9 | 10 | 400 | 1024 | Busy
10 | 11 | 800 | 2048 | Breakdown
11 | 12 | 1000 | 4096 | Diverter 2 Pos. A
12 | 13 | 2000 | 8192 | Diverter 2 Pos. B
13 | 14 | 4000 | 16384 | Diverter 3 Pos. A
14 | 15 | 8000 | 23768 | Diverter 3 Pos. B
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HR-LSP | 20 | In robot gripper | 20 | 10 | Empty capsule
2 |  |  | 21 | In lid opener | 21 | 12 | Start dosing
3 | [5 empty] | 13 | Start Cleaning
4 | 1 | HR-LSP Output | 30 | Dosing | 30 | 51 | Start transport HR-BM -> LAB
5 |  |  | 31 | In cup for extern | 31 | //52 | Start transport HR-BM(Hot meal) -> LAB
6 |  |  | 32 | Cup available | 32 | //53 | Start transport HR-LKC(Clinker) -> LAB
7 |  |  | 33 | Material waste | 33 | //54 | Start transport HR-BM(Cement) -> LAB
8 |  |  | 34 | Ready to fill external cup | 34 | 71 | Start transport LAB -> HR-BM
9 |  |  | 35 | Ready to fill external cup | 35 | //72 | Start transport LAB -> HR-BM(Hot meal)
10 |  |  | 36 | Transport Active | 36 | //73 | Start transport LAB -> HR-LKC(Clinker)
11 |  |  | 37 | Cup on transport belt | 37 | //74 | Start transport LAB -> HR-BM(Cement)
12 |  |  | 38 | Cup at output position | 38 | 14 | Cup from output to belt
13 |  |  | 60 | Start dosing | 60 | 15 | Cup from belt to output
14 |  |  | 70 | Wait for start 1 | 70
15 |  |  | 71 | Wait for start 2 | 71
16 |  |  | 72 | Wait for start 3 | 72
17 |  |  | 73 | Wait for start 4 | 73
18 |  |  | 74 | Wait for start 5 | 74
19 |  |  | 81 | Dosing wait | 81
20 |  |  | 82 | Dosing wait | 82
21 |  |  | 91 | Cup start | 91
22 |  |  | 92 | Cup start | 92
|  |  | 93 | To be cleaned | 93
|  |  | 94 | Move bottle | 94
|  |  | 96 | Wait till upper magazine is turned | 96
23 |  |  | 97 | Wait till lower magazine is turned | 97
24
25 | 2 | Tube-1 | 41 | Carrier arrived | 41
26 |  |  | //42 | Carrier arrived Hot meal | 42
27 |  |  | //43 | Carrier arrived Clinker | 43
28 |  |  | //44 | Carrier arrived Cement | 44
29 |  |  | 61 | Carrier in tube | 61
30 |  |  | //62 | Carrier Hot meal in tube | 62
31 |  |  | //63 | Carrier Clinker in tube | 63
32 |  |  | /64 | Carrier Cement in tube | 64
33 |  |  | 51 | Carrier arrived extern machine | 51
34 |  |  | //52 | Carrier arrived extern machine | 52
35

## HR-LA
Name | HR-LA | Receiving station |  |  |  |  | BACK TO MAIN
Unit No | 10 | HP-LA
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Wait for dosing or cleaning
3 | 4 | 10 | 16 | cup in manual input
4 | 5 | 20 | 32 | Cup in output position
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn reqested
7 | 10 | 400 | 1024 | Busy
8 | 11 | 800 | 2048 | Carrier in Lab
9 | 12 | 1000 | 4096 | Ready for cleaning
10 | 13 | 2000 | 8192 | Container Suction available
| 101 | Tube
1 | 0 | 1 | 1 | Ready
2 | 4 | 10 | 16 | Capsule correct inside
3 | 5 | 20 | 32 | Tube closed
4 | 8 | 100 | 256 | Syn is running
5 | 9 | 200 | 512 | Syn reqested
6 | 10 | 400 | 1024 | Busy
7 | 11 | 800 |  | Capsule wrong inside
8 | 12 | 1000 |  | Station locked
9 | 14 | 4000 |  | Passing sensor 1 (near Lab)
10 | 15 | 8000 |  | Passing sensor 2 (near field)
| 102 | Belt
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Cup in magazin input (cover locked)
3 | 4 | 10 | 16 | Sample in magazine available
4 | 5 | 20 | 32 | Scan sample allowed
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn reqested
7 | 10 | 400 | 1024 | Busy
| 103 | Magazine
1 | 0 | 1 | 1 | Ready
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn reqested
7 | 10 | 400 | 1024 | Busy
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HR-LA | 2 | Sample in gripper | 21 | 1 | Reset breakdown
|  |  | 3 | Dosing | 22 | 2 | Switch to manual mode
|  |  | 4 | output | 23 | 3 | Stop machine
|  |  | 7 | Cup cleaning | 26 | 4 | Start synchronisation
|  |  | 11 | Waste container | 30 | 5 | Switch to auto mode
|  |  | 31 | Cup in output pos | 31 | 13 | Cup from manual input to belt
|  |  | 52 | Passing sensor 01 | 52 | 14 | Cup from belt to manual input
|  |  | 53 | Passing sensor 02 | 53 | 15 | Cup from belt to output position
2 |  |  | 54 | Passing sensor 03 | 54 | 16 | Cup from output position to belt
4 | 1 | Tube | 1 | Sample arrived | 20 | 17 | Cup from LA to HP-MA
5 |  |  | 8 | In transport SS1 | 27 | 18 | Cup from HP-MA to LA
6 |  |  | 9 | In transport SS2 | 28 | 20 | Start dosing into cup
7 |  |  | 10 | In transport SS3 | 29 | 22 | Start take and empty capsule
25 | 2 | Belt | 5 | Manual input | 24 | 32 | Start take and empty capsule
26 |  |  | 6 | middle of belt | 25 | 38 | unlock magazine cover (Sample input in magzine)
27 |  |  | 20 | From Mag to HP-MA | 44 | 39 | Start from magazine to MA
28 |  |  | 91 | Insert in magazine | 200 | 41 | Cup from magazine to MA (cup from left side through the machine)

## HP-M
Name | HP-M | Milling Machine |  |  |  |  | BACK TO MAIN
Unit No | 4 | HP-M
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | //1 | 2 | 2 | Ready to take
3 | //2 | 4 | 4 | Ready to bring
4 | //3 | 8 | 8 | Input belt free
5 | //4 | 10 | 16 | Lack of material
7 | 8 | 100 | 256 | Syn is running
8 | 9 | 200 | 512 | Syn reqested
9 | 10 | 400 | 1024 | Busy
| 40 | Input1
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Open
3 | 4 | 10 | 16 | Closed
4 | 10 | 400 | 1024 | Busy
5 | 12 | 1000 | 4096 | Cup at input
| 41 | Transport1
1 | 0 | 1 | 1 | Ready
2 | //1 | 2 | 2 | Ready to take
3 | //2 | 4 | 4 | Ready to bring
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HP-M | 1 | Manual input | 10 | 10 | Open input
2 |  |  | 2 | Cup back to input | 11 | 11 | Close input
3 |  |  | 3 | Position belt | 12 | 12 | Cup from input to belt
4 |  |  | 4 | Position belt back | 13 | 13 | Cup from belt to input
5 |  |  | 5 | Cup filled | 14 | 14 | Gripper to manual input with cup
6 |  |  | 6 | Milling | 15 | 15 | Gripper to belt input with cup
7 |  |  | 7 | Dosing | 16 | 20 | Start sample from manual input
8 |  |  | 8 | Transport | 17 | 21 | Start sample from belt
9 |  |  | 9 | Ready to give out | 18 | 22 | Start cup from manual input
10 |  |  | 10 | Waste bin | 19 | 23 | Start cup from belt
11 |  |  | 11 | Splitt | 20 | 31 | Start Cleaning
12 |  |  | 50 | Sample Registration | 50
13

## Transport
Name | Transport Belt | [5 empty] | BACK TO MAIN
Unit No | 23 | Transport 1
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Breakdown
3 | 4 | 10 | 16 | Cup at HR-LSP
4 | 5 | 20 | 32 | Cup at cup cleaning
5 | 8 | 100 | 256 | Syn is running
6 | 9 | 200 | 512 | Syn is requested
7 | 10 | 400 | 1024 | Busy
8 | 12 | 1000 | 4096 | Bit12
9 | 13 | 2000 | 8192 | Bit13
10 | 14 | 4000 | 16384 | Bit14
11 | 15 | 8000 | 32768 | Bit15
| 24 | Transport 2
12 | 0 | 1 | 1 | Ready
13 | 3 | 8 | 8 | Breakdown
14 | 4 | 10 | 16 | Cup at output 1
15 | 5 | 20 | 32 | Cup at output 2
16 | 8 | 100 | 256 | Syn is running
17 | 9 | 200 | 512 | Syn is requested
18 | 10 | 400 | 1024 | Busy
19 | 12 | 1000 | 4096 | Output 1 closed
20 | 13 | 2000 | 8192 | Output 1 locked
21 | 14 | 4000 | 16384 | Output 2 closed
22 | 15 | 8000 | 32768 | Output 2 locked
| 25 | Transport 3 to Fusion Device 1
23 | 0 | 1 | 1 | Ready
24 | 3 | 8 | 8 | Breakdown
25 | 4 | 10 | 16 | Cup at fusion device 1
26 | 5 | 20 | 32 | Bit5
27 | 8 | 100 | 256 | Syn is running
28 | 9 | 200 | 512 | Syn is requested
29 | 10 | 400 | 1024 | Busy
30 | 12 | 1000 | 4096 | Fusion-1 Ready
31 | 13 | 2000 | 8192 | Bit13
32 | 14 | 4000 | 16384 | Bit14
33 | 15 | 8000 | 32768 | Bit15
| 26 | Transport 4 to Fusion Device 2
34 | 0 | 1 | 1 | Ready
35 | 3 | 8 | 8 | Breakdown
36 | 4 | 10 | 16 | Cup at fusion device 2
37 | 5 | 20 | 32 | Bit5
38 | 8 | 100 | 256 | Syn is running
39 | 9 | 200 | 512 | Syn is requested
40 | 10 | 400 | 1024 | Busy
41 | 12 | 1000 | 4096 | Fusion-2 Ready
42 | 13 | 2000 | 8192 | Bit13
43 | 14 | 4000 | 16384 | Bit14
44 | 15 | 8000 | 32768 | Bit15
| 27 | Cup input
45 | 0 | 1 | 1 | Ready
46 | 3 | 8 | 8 | Breakdown
47 | 4 | 10 | 16 | Cup at seperator
48 | 5 | 20 | 32 | Bit5
49 | 8 | 100 | 256 | Syn is running
50 | 9 | 200 | 512 | Syn is requested
51 | 10 | 400 | 1024 | Busy
52 | 12 | 1000 | 4096 | Bit12
53 | 13 | 2000 | 8192 | Bit13
54 | 14 | 4000 | 16384 | Bit14
55 | 15 | 8000 | 32768 | Bit15
| 28 | Average sample magazine
56 | 0 | 1 | 1 | Ready
57 | 3 | 8 | 8 | Breakdown
58 | 4 | 10 | 16 | Bottle in fill position
59 | 5 | 20 | 32 | Turntable in position
60 | 8 | 100 | 256 | Syn is running
61 | 9 | 200 | 512 | Syn is requested
62 | 10 | 400 | 1024 | Busy
63 | 12 | 1000 | 4096 | Door closed
64 | 13 | 2000 | 8192 | Door locked
65 | 14 | 4000 | 16384 | Bit14
66 | 15 | 8000 | 32768 | Bit15
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Transport 1 | 0 | Transport active | 10 | 10 | Full cup from HR-LSP to HP-M
2 |  |  | 1 | Cup cleaning active | 11 | 11 | Full cup from HR-LSP to average sample magazine Bottle 1
3 |  |  | 2 | Empty cup at HR-LSP | 12 | 12 | Full cup from HR-LSP to average sample magazine Bottle 2
4 |  |  | 3 | Sample at HP-M | 13 | 13 | Full cup from HR-LSP to average sample magazine Bottle 3
5 | [5 empty] | 14 | Full cup from HR-LSP to average sample magazine Bottle 4
6 | 1 | Transport 2 | 0 | Sample at Average sample magazine | 15 | 15 | Full cup from HR-LSP to cup output 1
7 |  |  | 1 | Sample dumped in bottle | 16 | 16 | Full cup from HR-LSP to cup output 2
8 |  |  | 2 | Cup at cup output 1 | 17 | 20 | Full cup from HP-M to Fusion device 1
9 |  |  | 3 | Cup at cup output 2 | 18 | 21 | Full cup from HP-M to Fusion device 2
10 |  |  | 90 | Empty cup at average sample magazine | 90 | 22 | Full cup from HP-M to cup output 1
11 | [5 empty] | 23 | Full cup from HP-M to cup output 2
12 | 2 | Transport 3 to Fusion Device 1 | 0 | Sample at Fusion device 1 | 20 | 30 | Empty cup from fusion device 1 to HR-LSP
13 |  |  | 90 | Empty cup at fusion device 1 | 91 | 31 | Empty cup from fusion device 2 to HR-LSP
14 | [5 empty] | 32 | Empty cup from average sample magazine to HR-LSP
15 | 3 | Transport 4 to Fusion Device 2 | 0 | Sample at Fusion device 2 | 21 | 33 | Empty cup from cup input to HR-LSP
16 |  |  | 90 | Empty cup at fusion device 2 | 92 | 40 | Insert empty cup in cup input
17 | [5 empty] | 41 | Unlock Cup output 1
18 | 4 | Cup input | 0 | Cup inserted in cup input | 25 | 42 | Unlock Cup output 2

## Rigaku XRF
Name | Rigaku XRF | [5 empty] | BACK TO MAIN
Unit No | 40
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 10 | 400 | 1024 | Busy
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | XRF-1 | 1 | Reg.Sample: Waiting | 1 | 5 | Reset error messages of XRF
2 |  |  | 2 | Reg.Sample OK | 2 | 10 | Register sample
3 |  |  | 3 | Reg.Sample: Error | 3 | 11 | Register sample (Prog. =1)
4 |  |  | 4 | Prep.Done: Waiting | 4 | 20 | Sample prepartion done
5 |  |  | 5 | Prep.Done: OK | 5 | 21 | Sample prepartion done (Prog. =1)
6 |  |  | 6 | Prep.Done: Error | 6 | 30 | Start analysis
7 |  |  | 7 | Unloaded: Waiting | 7 | 31 | Start analysis (Prog. =1)
8 |  |  | 8 | Unloaded: OK | 8
9 |  |  | 9 | Unloaded: Error | 9
10 |  |  | 10 | Analysis Start | 10
11 |  |  | 11 | Analysis Result | 11
12 |  |  | 12 | Analysis End | 12
13 |  |  | 13 | ERROR | 13
14 |  |  | 99 | Request analyse | 99

## HS-F 3000
Name | HS-F 3000
Unit No | 4
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Setup
3 | 4 | 10 | 16 | Semiautomatic
4 | 8 | 100 | 256 | Syn is running
5 | 9 | 200 | 512 | Syn is requested
6 | 10 | 400 | 1024 | Busy
Unit No | 41 | Manual input
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Setup
3 | 8 | 100 | 256 | Syn is running
4 | 9 | 200 | 512 | Syn is requested
5 | 10 | 400 | 1024 | Busy
Unit No | 42 | Input/Output transport
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Setup
3 | 8 | 100 | 256 | Syn is running
4 | 9 | 200 | 512 | Syn is requested
5 | 10 | 400 | 1024 | Busy |  |  | BACK TO MAIN
Unit No | 43 | Automatic input
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Setup
3 | 8 | 100 | 256 | Syn is running
4 | 9 | 200 | 512 | Syn is requested
5 | 10 | 400 | 1024 | Busy
Unit No | 44 | Tool magazine
1 | 0 | 1 | 1 | Ready
2 | 3 | 8 | 8 | Setup
3 | 8 | 100 | 256 | Syn is running
4 | 9 | 200 | 512 | Syn is requested
5 | 10 | 400 | 1024 | Busy
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | HSF-3000 | 1 | Clamping | 4 | 10 | Start Steel round sample
2 |  |  | 2 | Measuring diam. | 5 | 11 | Start Steel Oval sample
3 |  |  | 3 | Measuring height | 6 | 12 | Start Steel Double-Thickness sample
4 |  |  | 4 | Milling | 7 | 13 | Start Iron round sample
5 |  |  | 5 | Deburring | 8
6 |  |  | 90 | Register auto | 90
7 |  |  | 91 | Register semi | 91
8 |  |  | 92 | Register manu | 92
9 | 1 | Manual input | 1 | In transport | 1
10 |  |  | 2 | In machine | 2
11 |  |  | 3 | Input | 3
12 | 2 | Input/Output transport | 1 | At HSF | 12
13 |  |  | 2 | In transport | 13
14 |  |  | 3 | At output | 14
15 |  |  | 4 | At input | 15
16 | 3 | Robot input | 1 | Robot input | 9
17 |  |  | 2 | In transport | 10
18 |  |  | 3 | At HSF | 11
19 | 4 | Tool magazine

## AccuPyc Density Meter
Name | Density 1/1 | [5 empty] | BACK TO MAIN
Unit No | 13
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Density 1/1 | 1 | Get Ready | 1
2 |  |  | 2 | Register sample | 2
3 |  |  | 3 | Wait to Purg | 3
4 |  |  | 4 | Purging | 4
5 |  |  | 5 | Zero | 5
6 |  |  | 6 | Running | 6
7 |  |  | 7 | Results | 7
8 |  |  | 8 | Finshed | 8
9 |  |  | 9 | Ending... | 9
10 |  |  | 10 | Error | 10
11 |  |  | 11 | Err.Samp.Weight | 11

## OBLF OES
Name | OBLF OES | [5 empty] | BACK TO MAIN
Unit No | 34
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 2 | 4 | 4 | Request for output allowed
3 | 3 | 8 | 8 | Part of recal running
4 | 5 | 10 | 16 | Request for output allowed rejected
5 | 10 | 100 | 256 | Reprep request/ask
6 | 11 | 200 | 512 | Filing possible
7 | 12 | 400 | 1024 | MAG-session disabled
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | OES | 1 | Move to new position | 1 | 10 | Register
2 |  |  | 2 | Roboter | 2 | 11 | Disconnect
3 |  |  | 3 | Messung OK | 3 | 15 | ARE_YOU_THERE
4 |  |  | 4 | Messung NOK | 4 | 16 | Register BLF - Production
5 |  |  | 5 | Camera | 5 | 17 | Register BLF - Test
6 |  |  | 6 | Stored in Magazine | 6 | 18 | Register BLF - Recal
7 |  |  | 7 | Trash Can | 7 | 20 | Abfunkpunkt 1
8 |  |  | 8 | Labeler | 8 | 21 | Abfunkpunkt 2
9 |  |  | 9 | free | 9 | 22 | Abfunkpunkt 3
10 |  |  | 20 | Spark test point 1 | 20 | 23 | Abfunkpunkt 4
11 |  |  | 21 | Spark test point 2 | 21 | 24 | Abfunkpunkt 5
12 |  |  | 22 | Spark test point 3 | 22 | 25 | Abfunkpunkt 6
13 |  |  | 23 | Spark test point 4 | 23 | 26 | Abfunkpunkt 7
14 |  |  | 24 | Spark test point 5 | 24 | 27 | Abfunkpunkt 8
15 |  |  | 25 | Spark test point 6 | 25 | 28 | Abfunkpunkt 9
16 |  |  | 26 | Spark test point 7 | 26 | 29 | Abfunkpunkt 10
17 |  |  | 27 | Spark test point 8 | 27 | 40 | Delete Registration
18 |  |  | 28 | Spark test point 9 | 28 | 50 | Sample prepared
19 |  |  | 29 | Spark test point 10 | 29
20 |  |  | 30 | Input position 1 | 30
21 |  |  | 31 | Input position 2 | 31
22 |  |  | 40 | Output position 1 | 40
23 |  |  | 41 | Output position 2 | 41
24 |  |  | 50 | Collecting tank | 50
25 |  |  | 51 | Collecting tank | 51
26 |  |  | 60 | Park position 1 | 60
27 |  |  | 61 | Park position 2 | 61
28 |  |  | 94 | Move to new position | 62
29 |  |  | 95 | OK total measurement | 63
30 |  |  | 96 | NOK total measurement | 64
31 |  |  | 97 | Registration not successful | 65
32 |  |  | 98 | Registration successful | 66
33 |  |  | 99 | Test Startposition | 67
34 |  |  | 90 | Output position taken | 90
35 |  |  | 100 | Position Magazin | 100
36 |  |  | 101 | Position Magazin | 101

## SuperQ XRF
Name | SuperQ XRF | [5 empty] | BACK TO MAIN
Unit No | 30
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
2 | 10 | 400 | 1024 | Busy
Unit No | 31 | Manual input
1 | 0 | 1 | 1 | Ready
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | XRF | 1 | In Instrument | 1 | 1 | Analyse sample (only test)
2 |  |  | 2 | In Instrument | 2 | 2 | Remove sample (only test)
3 |  |  | 3 | analysis completed | 3 | 3 | Analyse sample (only test)
4 |  |  | 4 | Ready for output | 4 | 4 | Remove sample (only test)
5 |  |  | 5 | Output | 5 | 10 | Analyse sample
6 |  |  | 6 | Removed | 6 | 11 | Ok for output sample
7 |  |  | 7 | Back on start-position | 7 | 12 | Not ok for output sample
8 |  |  | 51 | Start | 8 | 13 | Remove sample
9 |  |  | 52 | Start manual | 9

## Sympatec
Name | Analytical Database | [5 empty] | BACK TO MAIN
Unit No | 7
Sl.No. | Status Bits (B) | Busy Code |  | Descriptions
|  | Value (Hex) | Value (Dec)
1 | 0 | 1 | 1 | Ready
Sl.No | Internal Unit |  | Positions in Unit |  |  | Commands
| No | Name | NO | Name | ID Index | Code | Name
1 | 0 | Analytical database |  |  |  | 101 | Request product
[6 empty] | 104 | Are you there